## SENSe
The SENSe control the current measurement range, energy counting/window, and the data acquisition sequence.

SCPI command	Description
SENSe
:AHOur
:RESet {<channel>}	Resets the amp-hour (Ah) measurement to zero
:CURRent
[:DC]
RANGe?	Returns selected DC current measurement range
:DLOG
:FUNCtion
:CURRent {<bool>}, {<channel>}	Enables/disables output current internal data logging
:POWer {<bool>}, {<channel>}	Enables/disables output power internal data logging
:VOLTage {<bool>}, {<channel>}	Enables/disables output voltage internal data logging
:PERiod {<time>}	Sets the sample period for internal data logging
:TIME {<time>}	Sets the sample duration for internal data logging
:WHOur
:RESet {<channel>}	Resets the watt-hour (Wh) measurement to zero

### SENSe:AHOur:RESet
Syntax	SENSe:AHOur:RESet {<channel>}
Description	Not implemented yet / Use this command to reset the amp-hour counter to zero of the specified channel. If channels are coupled (in series or parallel) or in tracking mode, both channel’s counters will be reset to zero.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2|CH3|CH4|CH5|CH6	–
Usage example	SENSE:AHO:RES
Related Commands	FETCh:WHOur? / INSTrument:COUPle:TRACking / OUTPut:TRACk[:STATe] / SENSe:AHOur:RESet / SYSTem:CHANnel:INFOrmation:AHOur:TOTal?

### SENSe:CURRent[:DC]:RANGe
Syntax	SENSe:CURRent[:DC]:RANGe?
Modules	DCP405 DCP405B
Description	The query command returns the maximum DC current that can be measured on the range that is set by auto-ranging circuit.
Usage example	SENS:CURR:RANG? / 0.05
Errors	-241,"Hardware missing"
Related Commands	SENSe:CURRent[:DC]:RANGe:AUTO / SYSTem:CHANnel:MODel?

### SENSe:DLOG:FUNCtion:CURRent
Syntax	SENSe:DLOG:FUNCtion:CURRent {<bool>}, {<channel>} / SENSe:DLOG:FUNCtion:CURRent? {<channel>}
Description	Use this command to enable or disable output current internal data logging on the specified channel.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	OFF
	<channel>	Discrete	CH1|CH2|CH3|CH4|CH5|CH6	–
Return	The query command returns the status (0 or 1) of the output current internal data logging enabled on the specified channel
Usage example	SENS:DLOG:FUNC:CURR ON, CH1
Related Commands	INIT:DLOG / SENSe:DLOG:FUNCtion:POWer / SENSe:DLOG:FUNCtion:VOLTage / SENSe:DLOG:PERiod / SENSe:DLOG:TIME

### SENSe:DLOG:FUNCtion:POWer
Syntax	SENSe:DLOG:FUNCtion:POWer {<bool>}, {<channel>} / SENSe:DLOG:FUNCtion:POWer? {<channel>}
Description	Use this command to enable or disable output power internal data logging on the specified channel.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	OFF
	<channel>	Discrete	CH1|CH2|CH3|CH4|CH5|CH6	–
Return	The query command returns the status (0 or 1) of the output power internal data logging enabled on the specified channel
Usage example	SENS:DLOG:FUNC:VOLT OFF, CH2
Related Commands	INIT:DLOG / SENSe:DLOG:FUNCtion:CURRent / SENSe:DLOG:FUNCtion:VOLTage / SENSe:DLOG:PERiod / SENSe:DLOG:TIME

### SENSe:DLOG:FUNCtion:VOLTage
Syntax	SENSe:DLOG:FUNCtion:VOLTage {<bool>}, {<channel>} / SENSe:DLOG:FUNCtion:VOLTage? {<channel>}
Description	Use this command to enable or disable output voltage internal data logging on the specified channel.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	OFF
	<channel>	Discrete	CH1|CH2|CH3|CH4|CH5|CH6	–
Return	The query command returns the status (0 or 1) of the output voltage internal data logging enabled on the specified channel
Usage example	SENS:DLOG:FUNC:VOLT OFF, CH2
Related Commands	INIT:DLOG / SENSe:DLOG:FUNCtion:CURRent / SENSe:DLOG:FUNCtion:POWer / SENSe:DLOG:PERiod / SENSe:DLOG:TIME

### SENSe:DLOG:PERiod
Syntax	SENSe:DLOG:PERiod {<time>} / SENSe:DLOG:PERiod?
Description	Use this command to set the sample period in seconds. For the sample period, the / minimum is 20 ms, and it has to be in integral of 20 ms, while the maximum is 120 s. The entered value will be rounded to the nearest 20 ms integral.
Parameters	Name	Type	Range	Default
	<time>	NR2	0.02 – 120	0.02
Return	The query command returns the sample period in seconds.
Usage example	SENS:DLOG:PER 1
Related Commands	SENSe:DLOG:TIME

### SENSe:DLOG:TIME
Syntax	SENSe:DLOG:TIME {<time>} / SENSe:DLOG:TIME?
Description	This command sets the sample duration, the entered value is in seconds. For sample duration, the maximum is 86 400 000 seconds or 24 000 hours or 1000 days (depending of the available space on the installed SD Card, see the MMEMory:INFOrmation? command).
Parameters	Name	Type	Range	Default
	<time>	NR1	1 – 86400000	60
Return	The query command returns the sample duration in seconds.
Usage example	SENS:DLOG:TIME 3600
Related Commands	MMEMory:INFOrmation? / SENSe:DLOG:PERiod

### SENSe:WHOur:RESet
Syntax	SENSe:WHOur:RESet {<channel>}
Description	Not implemented yet / Use this command to reset the watt-hour counter to zero of the specified channel. If channels are coupled (in series or parallel) or in tracking mode, both channel’s counters will be reset to zero.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2|CH3|CH4|CH5|CH6	–
Usage example	SENS:WHO:RES
Related Commands	FETCh:AHOur? / INSTrument:COUPle:TRACking / OUTPut:TRACk[:STATe] / SENSe:WHOur:RESet / SYSTem:CHANnel:INFOrmation:WHOur:TOTal