## FETCh
Fetch commands return measurement data that has been previously acquired. FETCh queries do not generate new measurements, but allow additional measurement calculations from the same acquired data.

SCPI command	Description
FETCh
:AHOur? {<channel>}	Returns the delivered energy in amp-hours
WHOur? {<channel>}	Returns the delivered energy in watt-hours

### FETCh:AHOur?
Syntax	FETCh:AHOur? {<channel>}
Description	Not implemented yet / Use this command to query delivered energy on the specified channel in amp-hours accumulated after last power-on or SENSe:AHOur:RESet command. / This value is measured independently of channel’s total delivered energy in amp-hours stored in non-volatile memory that can be queried using the SYSTem:CHANnel:INFOrmation:AHOur:TOTal? Command. / If channels are coupled (in series or parallel) or in tracking mode, this command will return a sum of delivered energy on both channels.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2|CH3|CH4|CH5|CH6	–
Usage example	FETC:AHO? CH1 / 1340.30
Related Commands	FETCh:WHOur? / INSTrument:COUPle:TRACking / OUTPut:TRACk[:STATe] / SENSe:AHOur:RESet / SYSTem:CHANnel:INFOrmation:AHOur:TOTal?

### FETCh:WHOur?
Syntax	FETCh:WHOur? {<channel>}
Description	Not implemented yet / Use this command to query delivered energy on the specified channel in watt-hours accumulated after last power-on or SENSe:AHOur:RESet command. / This value is measured independently of channel’s total delivered energy in watt-hours stored in non-volatile memory that can be queried using the SYSTem:CHANnel:INFOrmation:WHOur:TOTal? Command. / If channels are coupled (in series or parallel) or in tracking mode, this command will return a sum of delivered energy on both channels.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2|CH3|CH4|CH5|CH6	–
Usage example	FETC:WHO? CH1 / 100.30 / OUTP 0 / FETC:WHO:RES CH1 / FETC:WHO? CH1 / 0
Related Commands	FETCh:AHOur? / INSTrument:COUPle:TRACking / OUTPut:TRACk[:STATe] / SENSe:WHOur:RESet / SYSTem:CHANnel:INFOrmation:WHOur:TOTal?