## ROUTe
The ROUTe command subsystem controls switching operations for relay modules such as SMX46 and PREL6 and defines channel labels.

Please note that when list of channels separated with comma is specified no space is allowed. For example (@111,121) instead of (@111, 121) or (@111 , 121).

SCPI command	Description
ROUTe
:CLOSe {<chanlist>}	Closes (turns on) specified relay matrix channels
:LABel
:CHANnel {<label>},{<channel>}	Sets relay channel label
:COLumn {<slot>},{<column>},{<label>}	Sets relay matrix column label
:ROW {<slot>},{<row>},{<label>}	Sets relay matrix row label
:OPEN {<chanlist>}	Opens (turns off) specified relay matrix channels

### ROUTe:CLOSe
Syntax	ROUTe:CLOSe {<chanlist>} / ROUTe:CLOSe? {<chanlist>}
Modules	SMX PREL
Description	This command closes (turns it on) the relay matrix channels specified by <chanlist>. <chanlist> has the form (@nrc) where n is the module slot number, r is the channel row number and c is the column number.
Parameters	Name	Type	Range	Default
	<chanlist>	ChannelList	SMX: (@n03) and (@n11) to (@n46), PREL: (@n01) to (@n06) where n is the slot number	–
Return	Returns the current state of the channel(s) queried. The command returns "1" if channel(s) are closed (relay is energized) or returns "0" if channel(s) are open (relay is idle).
Usage example	Close (activate) four crosspoints (X1Y1, X2Y1, X3Y1, X4Y1) at once on the SMX46 inserted into slot 1: / ROUT:CLOS (@111,112,113,114) / Close all crosspoints in row 3 (X3Y1 to X3Y6) at once on the SMX46 inserted into slot 1: / ROUT:CLOS (@113:163) / Close power relay on SMX46 inserted in slot 2: / ROUT:CLOS (@203)
Errors	-200,"Execution error"
Related Commands	ROUTe:OPEN / SYSTem:RELay:CYCLes?

### ROUTe:LABel:CHANnel
Syntax	ROUTe:LABel:CHANnel {<channel>}, {<label>} / ROUTe:LABel:CHANnel? {<channel>}
Modules	PREL
Description	Use this command to define the channel label specified by <channel>.
Parameters	Name	Type	Range	Default
	<channel>	Channel	(@n01) to (@n06) where n is the slot number	–
	<label>	Quoted string	Max. 10 characters	–
Return	Returns the label of the channel queried as a quoted string.
Usage example	Set label “Heater” for the first channel on the PREL6 inserted into slot 2: / ROUT:CHAN:LAB (@201),"Heater"
Errors	-200,"Execution error"
Related Commands	ROUTe:CLOSe / ROUTe:OPEN

### ROUTe:LABel:COLumn
Syntax	ROUTe:LABel:COLumn {<slot>}, {<column>}, {<label>} / ROUTe:LABel:COLumn? {<slot>}, {<column>}
Modules	SMX
Description	Use this command to define the column label of the relay matrix specified by <slot> and <column>.
Parameters	Name	Type	Range	Default
	<slot>	NR1	1 – 3	–
	<column>	NR1	1 – 6	–
	<label>	Quoted string	Max. 5 characters	–
Return	Returns the label of the column queried on the specified slot as a quoted string.
Usage example	Set label “Sense” for the fifth column of the SMX46 inserted into slot 2: / ROUT:LAB:COL 2, 5, "Sense"
Errors	-200,"Execution error"
Related Commands	ROUTe:CLOSe / ROUTe:LABel:ROW / ROUTe:OPEN

### ROUTe:LABel:ROW
Syntax	ROUTe:LABel:ROW {<slot>}, {<row>}, {<label>} / ROUTe:LABel:ROW? {<slot>}, {<row>}
Modules	SMX
Description	Use this command to define the column label of the relay matrix specified by <slot> and <row>.
Parameters	Name	Type	Range	Default
	<slot>	NR1	1 – 3	–
	<row>	NR1	1 – 4	–
	<label>	Quoted string	Max. 5 characters	–
Return	Returns the label of the row queried on the specified slot as a quoted string.
Usage example	Set label “Bias” for the third row of the SMX46 inserted into slot 2: / ROUT:LAB:COL 2, 3, "Bias"
Errors	-200,"Execution error"
Related Commands	ROUTe:CLOSe / ROUTe:LABel:COLumn / ROUTe:OPEN

### ROUTe:OPEN
Syntax	ROUTe:OPEN {<chanlist>} / ROUTe:OPEN? {<chanlist>}
Modules	SMX PREL
Description	This command closes (turns it off) the relay matrix channels specified by <chanlist>. <chanlist> has the form (@nrc) where n is the module slot number, r is the channel row number and c is the column number. / The *RST command will open all channels of all modules.
Parameters	Name	Type	Range	Default
	<chanlist>	ChannelList	SMX: (@n03) and (@n11) to (@n46), PREL: (@n01) to (@n06) where n is the slot number	–
Return	Returns the current state of the channel(s) queried. The command returns "1" if channel(s) are open (relay is idle) or returns "0" if channel(s) are closed (relay is energized).
Usage example	Open (deactivate) two crosspoints (X1Y3 and X4Y2) at once on the SMX46 inserted into slot 1: / ROUT:OPEN (@131,124) / Open two sequence of crosspoints (X2Y2 to X4Y2 and X2Y4 to X4Y4, i.e. six crosspoints in total) at once on the SMX46 inserted into slot 1: / ROUT:OPEN (@122:124,142:144) / Open power relay on SMX46 inserted in slot 2: / ROUT:OPEN (@203)
Errors	-200,"Execution error"
Related Commands	*RST / ROUTe:CLOSe