## MMEMory
The MMEMory commands are used to store, read or delete file in the BB3’s SD card. It can also query SD card information. In addition it is used for storing and recalling values used by [SOURce[<n>]]:LIST subsystem. File and directory (folder) names cannot contain the following characters:\ / : * ? " < > |

SCPI command	Description
MMEMory
:CATalog [<directory>]	Returns a list of items in the specified directory (folder)
:LENgth [<directory>]	Returns the number of items in the specified directory
:CDIRectory {<directory>}	Changes the current directory
:CLOSe	Closes the file specified in NAME
:COPY {<source>}, {<destination>}	Copies <source> to <destination>
:DATE? {<filename>}	Returns date that the specified file was last saved
:DELete {<filename>}	Deletes an existing file
:DOWNload
:ABORt	Aborts current download session
:DATA {<block>}	Downloads data from the host computer
:FNAMe {<filename>}	Creates or opens the specified filename for download data
:SIZE {<filesize>}	Sets information about file size used for progress bar
:FEED	Sets data handle used to feed data into the file
:INFOrmation?	Returns used and free space
:LOAD
:LIST<n> {<filename>}	Loads stored LIST to the specified channel
:PROFile {<filename>}	Loads stored user profile
:LOCK {<password>}	Sets write protection
:MDIRectory {<directory>}	Makes a new directory
:MOVE {<source>}, {<destination>}	Moves or renames <source> to <destination>
:NAME {<filename>}	Sets the file name ro be opened or closed
:OPEN	Opens the file specified in NAME
:RDIRectory {<directory>}	Removes the specified directory
:STORe
:LIST<n> {<filename>}	Saves specified channel LIST
:PROFile {<filename>}	Saves specified user profile
:TIME? {<filename>}	Returns time that the specified file was last saved
:UNLock {<password>}	Clears write protection
:UPLoad? {<filename>}	Uploads data to the host computer

### MMEMory:CATalog
Syntax	MMEMory:CATalog? [<directory>]
Description	Returns the list of files and directories (folders) names, types and sizes in the current or specified directory. Number of items (files/directories) corresponds to the value returned by the MMEMory:CATalog:LENgth? command. / To read out the information in the root directory (folder), specify "\" (backslash) or "/". If <directory> is not set, this function is applied to the current directory. If directory=<path>, this function is applied to <current directory>\<path>. / Error occurs if the specified directory does not exist.
Parameters	Name	Type	Range	Default
	<directory>	Quoted string	Directory name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters	–
Return	Returns all files in the directory as list of comma delimited quoted strings of <filename>, <filetype> and <filesize>. The following file types are supported: / BIN – binary data / CSV – textual data (comma separated) / FOLD – directory (folder) / LIST – LIST program data / LOG – trace (display) or logged data (file extension .log) / PROF – user profile data (file extension .profile) / STAT – instrument (setting) state or user profiles (file extension .conf)
Usage example	MMEM:CAT? / "USER,FOLD,0","SCPI.PDF,BIN,1274844","SCH5B13A.PDF,BIN,296589", "Documents,FOLD,0","Lists,FOLD,0","Videos,FOLD,0", "profile0.profile,PROF,264" / MMEM:CAT? "USER" / "LST_2_3.CSV,BIN,88","FERY2.PDF,BIN,2443"
Errors	-250,"Mass storage error" / -251,"Missing mass storage" / -252,"Missing media" / -256,"File name not found"
Related Commands	MMEMory:CATalog:LENgth? / MMEMory:INFOrmation? / MMEMory:DATE? / MMEMory:TIME?

### MMEMory:CATalog:LENgth
Syntax	MMEMory:CATalog:LENgth? [<directory>]
Description	This command returns the number of items in the current or specified directory (folder). The result corresponds to the number of files returned by the MMEMory:CATalog? command.
Parameters	Name	Type	Range	Default
	<directory>	Quoted string	Directory name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters	–
Return	The query command returns number (NR1) of items (files and directories).
Usage example	MMEM:CAT:LEN? "USER" / 2
Errors	-256,"File name not found"
Related Commands	MMEMory:CATalog?

### MMEMory:CDIRectory
Syntax	MMEMory:CDIRectory {<directory>} / MMEMory:CDIRectory?
Description	Changes the current directory to the specified directory (folder). This directory must exist otherwise an error will be generated. / At *RST, this value is set to the root path.
Parameters	Name	Type	Range	Default
	<directory>	Quoted string	Directory name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters	–
Return	This query returns the full path of the current directory.
Usage example	MMEM:CDIR "TEST/Test folder2" / MMEM:CDIR? / "TEST/Test folder2"
Errors	-256,"File name not found"
Related Commands	*RST / MMEMory:CATalog? / MMEMory:MDIRectory / MMEMory:RDIRectory

### MMEMory:CLOSe
Syntax	MMEMory:CLOSe
Description	Not implemented yet / Use this command to close the selected file specified in NAME. An attempt to CLOSe a file that was not open generate an error. / This command is included only for compatibility with the SCPI standard and may not be used.
Usage example	MMEM:CLOS
Errors	-256,"File name not found"
Related Commands	*RST / MMEMory:NAME / MMEMory:OPEN

### MMEMory:COPY
Syntax	MMEMory:COPY {<source>}, {<destination>}
Description	Makes a copy of an existing file in the current directory. The file names must include any file extension. / If <destination> is a file name, the copy file is created in the current directory. When <destination> is a <path> (e.g. “test lists/test022”) the source file is duplicated in <current directory>/<path>.
Parameters	Name	Type	Range	Default
	<source>	Quoted string	Source file name, 1 to 255 characters	–
	<destination>	Quoted string	Copy file name or directory name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters	–
Usage example	MMEM:COPY "test.bin", "new2/test_new.bin"
Errors	-253,"Corrupt media" / -254,"Media full" / -256,"File name not found" / -258,"Media protected"
Related Commands	MMEMory:CATalog? / MMEMory:CDIRectory / MMEMory:LOCK / MMEMory:MDIRectory / MMEMory:MOVE

### MMEMory:DATE
Syntax	MMEMory:DATE? {<filename>}
Description	Returns the (year, month, day) that the specified file was last saved.
Parameters	Name	Type	Range	Default
	<filename>	Quoted string	File name or directory name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters	–
Return	Comma-separated numbers (NR1) of year, month, day formated as yyyy, mm, dd
Usage example	MMEM:DATE? "test.002" / 2017, 10, 1
Errors	-256,"File name not found"
Related Commands	MMEMory:TIME

### MMEMory:DELete
Syntax	MMEMory:DELete {<filename>}
Description	Use this command to delete a file in the current directory. If SD card is locked using the MMEMory:LOCK command, an error -258 will be generated.
Parameters	Name	Type	Range	Default
	<filename>	Quoted string	File name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters	–
Usage example	MMEM:DEL "test.002"
Errors	-256,"File name not found" / -258,"Media protected"
Related Commands	MMEMory:CATalog? / MMEMory:CDIRectory / MMEMory:RDIRectory / MMEMory:LOCK

### MMEMory:DOWNload:ABORt
Syntax	MMEMory:DOWNload:ABORt
Description	Use this command to abort currenlty running file transfer from the host initiated with the MMEMory:DOWNload:DATA command. If not active file transfer exists, the command will be ignored without generating any error.
Usage example	MMEM:ABOR
Related Commands	MMEMory:DOWNload:DATA

### MMEMory:DOWNload:DATA
Syntax	MMEMory:DOWNload:DATA {<block>}
Description	Downloads data from the host computer to a file in the SD card. This is a multiple steps process: / The filename must have been previously specified by MMEMory:DOWNload:FNAMe. / The data can be transferred in single or more blocks. Receiving of first block will erase all previously stored, and each consecutive block will be appended to the end / Download is finished when MMEMory:DOWNload:FNAMe with empty name was sent / Use MMEMory:INFOrmation? command first to check available space.
Parameters	Name	Type	Range	Default
	<block>	Data block	–	–
Usage example	Downloads text Hello world and store into the file "test file" in the current directory. Digit 2 denotes two digits of data length (11). / MMEM:DOWN:FNAM "test file" / MMEM:DOWN:DATA #211Hello world / MMEM:DOWN:FNAM ""
Errors	-253,"Corrupt media" / -254,"Media full" / -258,"Media protected"
Related Commands	MMEMory:DOWNload:FNAMe / MMEMory:INFOrmation? / MMEMory:LOCK

### MMEMory:DOWNload:FNAMe
Syntax	MMEMory:DOWNload:FNAMe {<filename>}
Description	Creates or opens the specified filename prior to writing data to that file with MMEMory:DOWNload:DATA.
Parameters	Name	Type	Range	Default
	<filename>	Quoted string	File name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters	–
Usage example	MMEMory:DOWNload:FNAMe "new_list.bin"
Errors	-258,"Media protected"
Related Commands	MMEMory:DOWNload:DATA

### MMEMory:DOWNload:SIZE
Syntax	MMEMory:DOWNload:SIZE {<filesize>}
Description	This command define filesize used by progress bar displayed on the local console during the file transfer. If filesize is not provided, progress bar will not be displayed.
Parameters	Name	Type	Range	Default
	<filesize>	NR1	0 to 2.147.483.648 (2 GiB) as limited by SD Card FAT	0
Usage example	MMEMory:DOWNload:SIZE 124000
Related Commands	MMEMory:DOWNload:DATA / MMEMory:DOWNload:FNAMe

### MMEMory:FEED
Syntax	MMEMory:FEED / MMEMory:FEED?
Description	Not implemented yet / Use this command to set data handle used to feed data into the file specified by MMEMory:NAME command. The data handle is fixed to HCOPy. This command is included only for compatibility with the SCPI standard, and may not be used.
Return	The command returns data handle name (that is fixed to HCOPy).
Usage example	MMEM:FEED? / HCOP
Related Commands	HCOPy:DESTination

### MMEMory:INFOrmation
Syntax	MMEMory:INFOrmation?
Description	Use this command to find out total amount of storage currently used and storage available on the SD card. The sum of that two amounts represents SD card capacity in bytes.
Return	The command returns used space and free space as two comma separated integers.
Usage example	MMEM:INFO? / 3932160,7732461568
Related Commands	MMEMory:DOWNload:DATA

### MMEMory:LOAD:LIST
Syntax	MMEMory:LOAD:LIST<n> {<filename>}
Description	Load stored LIST program from SD card location defined by <filename> to the channel defined with <n>.
Parameters	Name	Type	Range	Default
	<filename>	Quoted string	File name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters	–
Usage example	MMEM:LOAD:LIST1 "DC_DC conv testing.list" / MMEM:LOAD:LIST2 "DC_DC conv testing.list" / TRIG:SOUR BUS / INIT / *TRG
Errors	-200,"Execution error" / -256,"File name not found"
Related Commands	MMEMory:STORe:LIST[<n>] / [SOURce[<n>]]:LIST:COUNt / [SOURce[<n>]]:LIST:CURRent[:LEVel] / [SOURce[<n>]]:LIST:DWELl / [SOURce[<n>]]:LIST:VOLTage[:LEVel]

### MMEMory:LOAD:PROFile
Syntax	MMEMory:LOAD:PROFile {<filename>}
Description	This command loads data from user profile file to the user profile 0.
Parameters	Name	Type	Range	Default
	<filename>	Quoted string	File name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters	–
Usage example	MMEM:LOAD:PROF "old settings.profile"
Errors	-256,"File name not found"
Related Commands	*RCL / *SAV / MEMory:STATe:RECall:AUTO / MMEMory:STORe:PROFile

### MMEMory:LOCK
Syntax	MMEMory:LOCK {<password>} / MMEMory:LOCK?
Description	Use this command to enable write protection of SD card. All writing, deleting or modifying attempts on files or directories will generate an error if SD card is locked.
Parameters	Name	Type	Range	Default
	<password>	Quoted string	System password (4 to 16 characters)	–
Return	This query returns 0 if SD card is unlocked, or 1 if SD card is locked.
Usage example	MMEM:LOCK "test123" / MMEM:LOCK? / 1
Errors	122,"Invalid sys password"
Related Commands	MMEMory:UNLock / SYSTem:PASSword:NEW

### MMEMory:MDIRectory
Syntax	MMEMory:MDIRectory {<directory>}
Description	This command creates a new directory. If directory=<path>, this command creates a <current directory>/<path> directory.
Parameters	Name	Type	Range	Default
	<directory>	Quoted string	Directory name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters	–
Usage example	MMEM:MDIR "test folder"
Errors	-258,"Media protected"
Related Commands	MMEMory:CDIRectory / MMEMory:RDIRectory

### MMEMory:MOVE
Syntax	MMEMory:MOVE {<source>}, {<destination>}
Description	This command moves or renames an existing file. If destination is a file name, the source file is renamed to the new file name in the current directory. When destination=<path>, the source file is moved to <current directory>/<path>. / Error occurs if the source file does not exist or the destination file already exists.
Parameters	Name	Type	Range	Default
	<source>	Quoted string	Source file name, 1 to 255 characters	–
	<destination>	Quoted string	New file name or directory name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters	–
Usage example	Rename file: / MMEM:MOVE "old name","new name" / Move file from current directory to “/Test” directory: / MMEM:MOVE "new name","/Test/new name" / Move and rename file: / MMEM:MOVE "/Test/new name","/Documents/new doc"
Errors	-256,"File name not found" / -258,"Media protected"
Related Commands	MMEMory:CDIRectory / MMEMory:COPY

### MMEMory:MDIRectory
Syntax	MMEMory:MDIRectory {<directory>}
Description	This command creates a new directory. If directory=<path>, this command creates a <current directory>/<path> directory.
Parameters	Name	Type	Range	Default
	<directory>	Quoted string	Directory name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters	–
Usage example	MMEM:MDIR "test folder"
Errors	-258,"Media protected"
Related Commands	MMEMory:CDIRectory / MMEMory:RDIRectory

### MMEMory:NAME
Syntax	MMEMory:NAME {<filename>} / MMEMory:NAME?
Description	Not implemented yet / This command specifies the name of the file used by MMEMory:OPEN and MMEMory:CLOSe commands.
Parameters	Name	Type	Range	Default
	<filename>	Quoted string	File name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters	–
Return	This query returns the name of the file that can be used by MMEMory:OPEN and MMEMory:CLOSe commands.
Usage example	MMEM:NAME "screenimg.png" / MMEM:NAME? / Screenimg.png
Errors	-257,"File name error"
Related Commands	MMEMory:CLOSe / MMEMory:OPEN

### MMEMory:OPEN
Syntax	MMEMory:OPEN
Description	Not implemented yet / Use this command to open the file specified in NAME. / This command is included only for compatibility with the SCPI standard and may not be used
Usage example	MMEM:OPEN
Errors	-256,"File name not found"
Related Commands	*RST / MMEMory:CLOSe / MMEMory:NAME

### MMEMory:STORe:LIST
Syntax	MMEMory:STORe:LIST<n> {<filename>}
Description	Store LIST program of channel defined with <n> to SD card location defined by <filename>. Stored file type will be TRAC. Default file extension is .list
Parameters	Name	Type	Range	Default
	<filename>	Quoted string	File name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters	–
Usage example	MMEM:STOR:LIST1 "DC_DC conv testing.list"
Errors	-258,"Media protected"
Related Commands	MMEMory:LOAD:STATe / [SOURce[<n>]]:LIST:COUNt / [SOURce[<n>]]:LIST:CURRent[:LEVel] / [SOURce[<n>]]:LIST:DWELl / [SOURce[<n>]]:LIST:VOLTage[:LEVel]

### MMEMory:STORe:PROFile
Syntax	MMEMory:STORe:PROFile {<filename>}
Description	This command stores the BB3 state (i.e. user profile 0) to SD card location defined by <filename>. Stored file type will be STAT. Default file extension is .conf
Parameters	Name	Type	Range	Default
	<filename>	Quoted string	File name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters	–
Usage example	MMEM:STOR:PROF "Both channels 5V_3A.profile"
Errors	-258,"Media protected"
Related Commands	*RCL / *SAV / MEMory:STATe:RECall:AUTO / MMEMory:LOAD:PROFile

### MMEMory:TIME
Syntax	MMEMory:TIME? {<filename>}
Description	Returns the (hours, minute, seconds) that the specified file was last saved.
Parameters	Name	Type	Range	Default
	<filename>	Quoted string	File name or directory name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters	–
Return	Comma-separated numbers (NR1) of hours, minute, seconds formatted as hh, mm, ss
Usage example	MMEM:TIME? "test.002" / 22, 10, 14
Errors	-256,"File name not found"
Related Commands	MMEMory:DATE

### MMEMory:UNLock
Syntax	MMEMory:UNLock {<password>}
Description	Disable write protection of SD card activated by the MMEMory:LOCK command.
Parameters	Name	Type	Range	Default
	<password>	Quoted string	System password (4 to 16 characters)	–
Usage example	MMEM:LOCK? / 1 / MMEM:UNL "test123" / MMEM:LOCK? / 0
Errors	122,"Invalid sys password"
Related Commands	MMEMory:LOCK / SYSTem:PASSword:NEW

### MMEMory:UPLoad
Syntax	MMEMory:UPLoad? {<filename>}
Description	This command uploads the contents of a file from the instrument to the host computer. / The format for <file> is "[<path>]<file_name>", where <path> must be an absolute folder path. If <path> is omitted, the folder specified by the MMEMory:CDIRectory command is used. Absolute paths begin with a "\" or "/" and start at the root folder of SD card.
Parameters	Name	Type	Range	Default
	<filename>	Quoted string	File name or directory name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters	–
Return	The query returns the file contents are returned as an IEEE 488.2 definite-length block.
Usage example	MMEM:UPL? "test file" / #211Hello world
Errors	-257,"File name error"
Related Commands	MMEMory:DOWNload:DATA