## MEASure
Measure commands return back the output voltage, current, power or temperature. They trigger the acquisition of new data before returning the reading. Measurements are performed by digitizing the instantaneous output voltage, current or temperature. Output power is calculated as product of measured voltage and current. Keyword [:DC] is optional since all measurement are by default of the DC level of the signal.

SCPI command	Description
MEASure
[:SCALar]
:CURRent
[:DC]? [<channel>]	Takes a measurement; returns the average current
:POWer
[:DC]? [<channel>]	Takes a measurement; returns the average power
:TEMPerature
[:DC]? {<sensor>}	Takes a measurement; returns the average temperature
[:VOLTage]
[:DC]? [<channel>]	Takes a measurement; returns the average voltage

### MEASure[:SCALar]:CURRent[:DC]
Syntax	MEASure[:SCALar]:CURRent[:DC]? [<channel>]
Description	Query the current measured across the current sense resistor inside the power module.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2|CH3|CH4|CH5|CH6	–
Return	Returns the average output current in amperes as decimal number (NR2).
Usage example	Measure current on the currently selected channel (CH1) and CH2: / MEAS:CURR?;:MEAS:CURR? CH2 / 1.23;0.12
Related Commands	INSTrument:COUPle:TRACking

### MEASure[:SCALar]:POWer[:DC]
Syntax	MEASure[:SCALar]:POWer[:DC]? [<channel>]
Description	Query the output power calculated as product of measured voltage and current value.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2|CH3|CH4|CH5|CH6	n/a
Return	Returns the average output power in watts as decimal number (NR2).
Usage example	MEAS:POW? CH2 / 80.44
Related Commands

### MEASure[:SCALar]:TEMPerature[:THERmistor][:DC]
Syntax	MEASure[:SCALar]:TEMPerature[:THERmistor][:DC]? {<sensor>}
Description	Query the temperature measured using the specified temperature sensors as follows: / AUX – auxiliary temperature sensor (built-in fan controller) used to monitor temperature inside BB3 chassis / CH1, CH2, CH3, CH4, CH5, CH6 – temperature sensor connected to channel power module used to regulate fan speed / All installed sensors are periodically tested. When sensor does not passed the test, programmed output current is automatically limited to 2 A. If load that draws more current then 2 A output current will be set to zero.
Parameters	Name	Type	Range	Default
	<sensor>	Discrete	AUX|CH1|CH2|CH3|CH4|CH5|CH6	AUX
Return	Returns the average temperature value in degrees Celsius (oC) as decimal number (NR2). If self-test detect selected temperature sensor failure or sensor is not installed a -240 or -241 error will be generated.
Usage example	MEAS:TEMP? CH2 / 49
Errors	-240,"Hardware error" / -241,"Hardware missing"
Related Commands	DIAGnostic[:INFOrmation]:FAN? / SYSTem:CPU:OPTion? / SYSTem:TEMPerature:PROTection[:HIGH][:LEVel] / SYSTem:TEMPerature:PROTection[:HIGH]:DELay[:TIME] / SYSTem:TEMPerature:PROTection[:HIGH]:STATe / SYSTem:TEMPerature:PROTection[:HIGH]:TRIPped?

### MEASure[:SCALar][:VOLTage][:DC]
Syntax	MEASure[:SCALar][:VOLTage][:DC]? [<channel>]
Description	Query the voltage measured at the sense terminals of the selected channel.
Parameters	Name	Type	Range	Default
	<channel>	Discrete	CH1|CH2|CH3|CH4|CH5|CH6	n/a
Return	Returns the average output voltage in volts as decimal number (NR2).
Usage example	MEAS:VOLT? CH1 / 43.25 / Query voltage of the channel 2 that is currently selected: / INST CH2 / MEAS? / 12.40
Related Commands	INSTrument:COUPle:TRACking