## ROUTe
The ROUTe command subsystem controls switching operations for relay modules such as SMX46 and PREL6 and defines channel labels.

Please note that when list of channels separated with comma is specified no space is allowed. For example (@111,121) instead of (@111, 121) or (@111 , 121).

SCPI command	Description
ROUTe
:CHANnel
:LABel {<label>},{<chanlist>}	Sets labels for relay matrix channels
:CLOSe {<chanlist>}	Closes (turns on) specified relay matrix channels
:OPEN {<chanlist>}	Opens (turns off) specified relay matrix channels

### ROUTe:CLOSe
Syntax	ROUTe:CLOSe {<chanlist>} / ROUTe:CLOSe? {<chanlist>}
Modules	SMX PREL
Description	This command closes (turns it on) the relay matrix channels specified by <chanlist>. <chanlist> has the form (@nrc) where n is module slot number, r is channel row number and c is column number.
Parameters	Name	Type	Range	Default
	<chanlist>	ChannelList	SMX: (@n03) and (@n11) to (@n46), PREL: (@n01) to (@n05) where n is a slot number	–
Return	Returns the current state of the channel(s) queried. The command returns "1" if channel(s) are closed (relay is energized) or returns "0" if channel(s) are open (relay is idle).
Usage example	Close (activate) four crosspoints (X1Y1, X2Y1, X3Y1, X4Y1) at once on the SMX46 inserted into slot 1: / ROUT:CLOS (@111,112,113,114) / Close all crosspoints in row 3 (X3Y1 to X3Y6) at once on the SMX46 inserted into slot 1: / ROUT:CLOS (@113:163) / Close power relay on SMX46 inserted in slot 2: / ROUT:CLOS (@203)
Errors	-200,"Execution error"
Related Commands	ROUTe:OPEN / SYSTem:RELay:CYCLes?

### ROUTe:CHANnel:LABel
Syntax	ROUTe:CHANnel:LABel: {label}, {<chanlist>} / ROUTe:CHANnel:LABel? {<chanlist>}
Modules	SMX PREL
Description	Use this command to define the column label of the relay matrix channels specified by <row>.
Parameters	Name	Type	Range	Default
	<label>	Quoted string	Max. 5 characters	–
	<chanlist>	ChannelList	SMX: (@n03) and (@n11) to (@n46), PREL: (@n01) to (@n05) where n is a slot number	–
Return	Returns the label of the channel queried.
Usage example	Set label “AC in” for the switch matrix row 1 on the SMX46 inserted into slot 1: / ROUT:CHAN:LAB "AC in",(@111) / Set label “Master” for the power relay on the SMX46 inserted into slot 1: / ROUT:CHAN:LAB "Master",(@103)
Errors	-200,"Execution error"
Related Commands	ROUTe:CLOSe / ROUTe:OPEN

### ROUTe:OPEN
Syntax	ROUTe:OPEN {<chanlist>} / ROUTe:OPEN? {<chanlist>}
Modules	SMX PREL
Description	This command closes (turns it off) the relay matrix channels specified by <chanlist>. <chanlist> has the form (@nrc) where n is module slot number, r is channel row number and c is column number. / The *RST command will open all channels of all modules.
Parameters	Name	Type	Range	Default
	<chanlist>	ChannelList	SMX: (@n03) and (@n11) to (@n46), PREL: (@n01) to (@n05) where n is a slot number	–
Return	Returns the current state of the channel(s) queried. The command returns "1" if channel(s) are open (relay is idle) or returns "0" if channel(s) are closed (relay is energized).
Usage example	Open (deactivate) two crosspoints (X1Y3 and X4Y2) at once on the SMX46 inserted into slot 1: / ROUT:OPEN (@131,124) / Open two sequence of crosspoints (X2Y2 to X4Y2 and X2Y4 to X4Y4, i.e. six crosspoints in total) at once on the SMX46 inserted into slot 1: / ROUT:OPEN (@122:124,142:144) / Open power relay on SMX46 inserted in slot 2: / ROUT:OPEN (@203)
Errors	-200,"Execution error"
Related Commands	*RST / ROUTe:CLOSe