## SYSTem
System commands control system functions that are not directly related to output control, measurement, or status functions.

SCPI command	Description
SYSTem
:BEEPer[:IMMediate]	Issues a single beep immediately
:KEY
:STATe {<bool>}	Enables click tone for local control
:STATe {<bool>}	Enables beeper function
:CAPability?	Returns an <instrument_specifier>
:CHANnel
:COLor {<channel>}, {<color>}	Sets channel color
[:COUNt]?	Returns the number of output channels
:INFOrmation
:CURRent?	Returns output current capability
:ONTime
LAST?	Returns time passed after last output enable
TOTal?	Returns channel’s total active time
:POWer?	Returns output power capability
:VOLTage?	Returns output voltage capability
:LABel {<channel>}, {<label>}	Sets channel label name
:MODel?	Returns the channel model and version name
:OPTion? [{<channel>}]	Returns names of all channel resources
:PIN
:LABel {<channel>}, {<pin>}, {<label>}	Sets module pin label name
:SLOT?	Returns the channel slot number
:VERSion?	Returns the channel version number
:COMMunicate
:ENABle {<bool>}, {<interface>}	Enables the remote interface
:ETHernet
:ADDRess {<ip_address>}	Sets the static LAN (IP) address
:CONTrol?	Returns the control connection port
:DHCP {<bool>}	Enables the use of the Dynamic Host Configuration Protocol (DHCP)
:DNS <ip_address>	Sets the IP address of the DNS server.
:GATEway {<ip_address>}	Sets the IP address of the default gateway
:HOSTname {<name>}	Sets the Ethernet communication host name
:MAC?	Returns the MAC address
:PORT {<number>}	Sets the port number
:SMASk {<mask>}	Sets the static subnet mask
:NTP {<server>}	Sets NTP service server address
:FREQuency	Sets the frequency of connection to the NTP service
:MQTT
:SETTings {<address>, <port>, <user>, <password>, <period>}	Sets MQTT connection parameters
:STATe?	Returns MQTT connection status
:USB
:CLAss {<usb_class>}	Sets USB class
:MODE {<usb_mode>}	Sets USB mode
:RLSTate {<state>}	Places the instrument in remote or local mode
:CPU
:FIRMware?	Returns BB3 firmware version
:INFOrmation
:ONtime
LAST?	Returns time passed after last power on
TOTal?	Returns the BB3 total active time
:MODel?	Returns the control board model name
:SNO?	Returns the BB3 serial number
:VERSion?	Returns the BB3 version number
:DATE {<year>}, {<month>}, {<day>}	Sets the date of the system clock
DELay {<time>}	Sets a pause in SCPI execution
:DIGital
:INPut:DATA? {<pin>}	Reads the state of the digital port pins
:OUTPut
:DATA {<pin>}, {<state>}	Sets the state of the digital port pins
:PWM
:DUTY {<pin>}, {<duty>}	Sets square wave generator duty cycle
:FREQuency {<pin>}, {<frequency>}	Sets square wave generator frequency
:PIN<n>
:FUNCtion {<function>}	Sets the selected pin’s function
:POLarity {<polarity>}	Sets the selected pin’s polarity
:ERRor
[:NEXT]?	Queries and clears errors from the error queue
:COUNt?	Queries the error/event queue for the number of unread items
:FAN
:SPEed	Returns speed of the cooling fan
:STATus	Returns status of the cooling fan
:FORMat
:DATE	Sets format for displaying date
:TIME	Sets 12h or 24h clock format
:INHibit?	Queries system inhibit state
:KEY
:KLOCk	Disables front panel [lock/unlock] icon
:LOCal	Places the BB3 in the local mode
:MEASure
[:SCALar]
:TEMPerature
[:DC]? {<sensor>}	Takes a measurement; returns the average temperature
[:VOLTage]
[:DC]? {<device>}	Takes a voltage measurement of the RTC battery
:PASSword
:CALibration
:RESet	Resets the calibration password to initial value
:FPANel
:RESet	Resets the front panel lock password to initial value
:NEW {<old>}, {<new>}	Changes system password
:PON:OUTPut:DISable {<bool>}	Sets output state on power up
:POWer {<bool>}	Enters the BB3 into the standby mode
:PROTection:TRIP {<bool>}	Enters the BB3 into the standby mode in case of protection trip
:RELay:CYCLes? {<chanlist>}	Returns the cycle count on the specified channels
:REMote	Places the BB3 in the remote mode
:REStart	Initiate low-level (hardware) reset
:RWLock	Places the BB3 in the remote mode and disables front panel [lock/unlock] icon
:SLOT
[:COUNt]?	Returns the number of available slots
:COLor {<slot>}, {<color>}	Sets module color
:LABel {<slot>}, {<label>}	Sets module label name
:MODel? {<slot>}	Returns the module model name
:SNO? {<slot>}	Returns the module serial number
:STATe {<slot>}, {<bool>}	Enables/disables module on selected slot
:VERSion? {<slot>}	Returns module version number
:TEMPerature
:PROTection
[:HIGH]
[:LEVel] {<temperature>} [, <sensor>]	Sets the OTP value
:CLEar {<sensor>}	Clears the latched protection status of the over-temperature protection (OTP)
:DELay
[:TIME] {<delay>} [, <sensor>]	Sets the OTP programming delay
:STATe {<bool>} [, <sensor>]	Enables/disables OTP on the selected temperature sensor
:TRIPped? [<sensor>]	Returns status of OTP activation
:TIME {<hours>}, {<minutes>}, {<seconds>}	Sets the time of the system clock
:DTS {rules}	Defines daylight saving time (DST) rules
:ZONE {zone}	Defines time zone
:VERSion?	Returns the SCPI version number

### SYSTem:BEEPer
Syntax	SYSTem:BEEPer[:IMMediate]
Description	This command issues a single beep immediately.
Usage example	SYST:BEEP

### SYSTem:BEEPer:KEY:STATe
Syntax	SYSTem:BEEPer:KEY:STATe {<bool>} / SYSTem:BEEPer:KEY:STATe?
Description	Use this command to enable or disable generation of audible “click” sound when front panel option is selected.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	OFF
Usage example	SYST:BEEP:KEY:STAT ON
Related Commands	SYSTem:BEEPer[:IMMediate] / SYSTem:BEEPer:STATe

### SYSTem:BEEPer:STATe
Syntax	SYSTem:BEEPer:STATe {<bool>} / SYSTem:BEEPer:STATe?
Description	When the beeper is enabled, the BB3 generates audible sound in any of the following situations: / the power is turns on or off (see SYSTem:POWer), / when error occurs during front panel operation or remote operation (see Section 7 for the list of error messages), / self-test is failed and / any of the protection function is “tripped”
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	ON
Usage example	SYST:BEEP:STAT ON / SYST:BEEP:STAT? / 1
Related Commands	SYSTem:BEEPer[:IMMediate] / SYSTem:POWer

### SYSTem:CAPability?
Syntax	SYSTem:CAPability?
Description	This query returns the BB3’s capabilities and outputs the appropriate specifiers. / See also SCPI Volume 4: Section 7.1, 1.4.1, 7.2.1, 7.2.2, and 7.2.3
Usage example	SYST:CAP? / DCPSUPPLY WITH (MEASURE|MULTIPLE|TRIGGER)

### SYSTem:CHANnel:COLor
Syntax	SYSTem:CHANnel:COLor {<channel>}, {<color>} / SYSTem:CHANnel:COLor? [<channel>]
Modules	DCP DCM
Description	Use this command to set channel header and frame color as defined with <color> value between 0 and 24. Use 0 to reset channel color to the default value. Default color number coresponds to the channel number. Available colors and its RGB values is listed below. / This command is only applicable to modules that have discrete channel numbers (CH1, CH2 ...). If used on other modules, e.g. SMX or PREL, it may generate -241,"Hardware missing" execution error.
	1 #d4c36b			2 #9fdd9a			3 #73bdd8		4 #d971ca
	5 #dbafea			6 #adaaf9			7 #f44336		8 #e91e63
	9 #9c27b0			10 #673ab7			11 #3f51b5		12 #2196f3
	13 #03a9f4			14 #00bcd4			15 #009688		16 #4caf50
	17 #8bc34a			18 #cddc39			19 #ffeb3b		20 #ffc107
	21 #ff9800			22 #ff5722			23 #795548		24 #607d8b

Parameters	Name		Type			Range			Default
	<channel>		Channel			CH1|CH2|CH3|CH4|CH5|CH6			–
	<color>		NR1			0 – 24			0
Return	Query the color number of the specified channel.
Errors	-222,"Data out of range" / -241,"Hardware missing"
Usage example	SYST:CHAN:COL CH1,2
Related Commands	SYSTem:SLOT:COLor

### SYSTem:CHANnel[:COUNt]?
Syntax	SYSTem:CHANnel[:COUNt]?
Description	This query returns the number of output channels in a mainframe.
Usage example	SYST:CHAN? / 2
Related Commands	INSTrument[:SELect] / INSTrument:NSELect

### SYSTem:CHANnel:INFOrmation:CURRent?
Syntax	SYSTem:CHANnel:INFOrmation:CURRent? [<channel>]
Description	Use this query to get currently selected channel output current capability.
Parameters	Name	Type	Range	Default
	<channel>	Channel	CH1|CH2|CH3|CH4|CH5|CH6	–
Usage example	SYST:CHAN:INFO:CURR? CH2 / 5.00

### SYSTem:CHANnel:INFOrmation:ONTime:LAST?
Syntax	SYSTem:CHANnel:INFOrmation:ONTime:LAST? [<channel>]
Description	This query returns time passed after last activation of the currently selected channel. Resolution is 1 minute and this information is stored every 10 minutes in non-volatile memory. Therefore it’s possible that up to 10 minutes is lost after restart caused with power outage or system reset.
Parameters	Name	Type	Range	Default
	<channel>	Channel	CH1|CH2|CH3|CH4|CH5|CH6	–
Usage example	SYST:CHAN:INFO:ONT:LAST? / 15m

### SYSTem:CHANnel:INFOrmation:ONTime:TOTal?
Syntax	SYSTem:CHANnel:INFOrmation:ONTime:TOTal? [<channel>]
Description	This query returns total active time of the currently selected channel. Resolution is 1 minute and this information is stored every 10 minutes in non-volatile memory. Therefore it’s possible that up to 10 minutes is lost after restart caused with power outage or system reset.
Parameters	Name	Type	Range	Default
	<channel>	Channel	CH1|CH2|CH3|CH4|CH5|CH6	–
Usage example	SYST:CHAN:INFO:ONT:TOT? / 1h 45m

### SYSTem:CHANnel:INFOrmation:POWer?
Syntax	SYSTem:CHANnel:INFOrmation:POWer? [<channel>]
Description	Use this query to get currently selected channel output power capability.
Parameters	Name	Type	Range	Default
	<channel>	Channel	CH1|CH2|CH3|CH4|CH5|CH6	–
Usage example	SYST:CHAN:INFO:POW? / 160.00

### SYSTem:CHANnel:INFOrmation:VOLTage?
Syntax	SYSTem:CHANnel:INFOrmation:VOLTage? [<channel>]
Description	Use this query to get currently selected channel output voltage capability.
Parameters	Name	Type	Range	Default
	<channel>	Channel	CH1|CH2|CH3|CH4|CH5|CH6	–
Usage example	SYST:CHAN:INFO:VOLT? / 40.00

### SYSTem:CHANnel:LABel
Syntax	SYSTem:CHANnel:LABel {<channel>}, {<label>} / SYSTem:CHANnel:LABel? {<channel>}
Description	Use this command to define the label of selected <channel>.
Parameters	Name	Type	Range	Default
	<channel>	Channel	CH1|CH2|CH3|CH4|CH5|CH6	–
	<label>	Quoted string	1 – 10 characters	–
Return	Returns the label of the channel queried as a quoted string.
Errors	-241,"Hardware missing"
Usage example	SYST:CHAN:LAB CH1,"Heater"
Related Commands	SYSTem:CHANnel:PIN:LABel / SYSTem:SLOT:LABel

### SYSTem:CHANnel:MODel?
Syntax	SYSTem:CHANnel:MODel? [<channel>]
Description	This query returns the model name and version of the specified channel.
Parameters	Name	Type	Range	Default
	<channel>	Channel	CH1|CH2|CH3|CH4|CH5|CH6	–
Usage example	SYST:CHAN:MOD? / "DCP405"
Related Commands	SYSTem:CHANnel:VERSion? / SYSTem:CPU:MODel?

### SYSTem:CHANnel:OPTion?
Syntax	SYSTem:CHANnel:OPTion? [<channel>]
Description	This query returns names of all selected channel resources that can be controlled by firmware. Depending of the board model (see the SYSTem:CHANnel:MODel? query) various combination of the following features can be returned: / Volt – program the output voltage while channel is in the CV mode of operation (see [SOURce[<n>]]:VOLTage and APPLy commands) / Current – program the output current while channel is in the CC mode of operation (see [SOURce[<n>]]:CURRent and APPLy commands) / Power – set max. allowed output power regardless of the channel mode of operation (see [SOURce[<n>]]:POWer:LIMit) / OE – set channel power output (see OUTPut[:STATe]) / Dprog – control down-programmer circuit (see OUTPut:DPRog) / Rprog – control output voltage programming source (see [SOURce[<n>]]:VOLTage:PROGram[:SOURce]) / Coupled – power outputs supports advanced coupling, i.e. in series and parallel (see
Parameters	Name	Type	Range	Default
	<channel>	Channel	CH1|CH2|CH3|CH4|CH5|CH6	–
Usage example	Reported resources for channel 1 on the DCP405 module: / SYST:CHAN:OPT? CH1 / "Volt", "Current", "Power", "OE", "DProg", "Rprog", "Coupled" / Reported resources for channel 3 on the DCM220 module: / SYST:CHAN:OPT? CH1 / "Volt", "Current", "Power", "OE"
Related Commands	APPLy / OUTPut:DPRog / OUTPut:MODE? / OUTPut[:STATe] / OUTPut:TRACk[:STATe] {<chanlist>} / [SOURce[<n>]]:CURRent / [SOURce[<n>]]:POWer:LIMit / [SOURce[<n>]]:VOLTage / [SOURce[<n>]]:VOLTage:PROGram[:SOURce]) / SYSTem:CHANnel:MODel?

### SYSTem:CHANnel:PIN:LABel
Syntax	SYSTem:CHANnel:PIN:LABel {<channel>}, {<pin>}, {<label>} / SYSTem:CHANnel:PIN:LABel? {<channel>}, {<pin>}
Modules	MIO
Description	Use this command to define the pin label specified by <label> on selected <channel>.
Parameters	Name	Type	Range	Default
	<channel>	NR1	1 – 3	–
	<pin>	NR1	1 – 8	–
	<label>	Quoted string	1 – 5 characters	–
Return	Returns the label of the pin queried as a quoted string.
Errors	-241,"Hardware missing"
Usage example	SYST:CHAN:PIN:LAB 1,"In_1"
Related Commands	SYSTem:CHANnel:LABel / SYSTem:SLOT:LABel

### SYSTem:CHANnel:SLOT?
Syntax	SYSTem:CHANnel:SLOT? [<channel>]
Description	This query returns the DIB backplane slot number of the specified channel.
Parameters	Name	Type	Range	Default
	<channel>	Channel	CH1|CH2|CH3|CH4|CH5|CH6	–
Usage example	SYST:CHAN:SLOT? CH4 / 3
Related Commands	SYSTem:CHANnel:SNO? / SYSTem:CHANnel:VERSion?

### SYSTem:CHANnel:VERSion?
Syntax	SYSTem:CHANnel:VERSion? [<channel>]
Description	This query returns the version number of the specified channel.
Parameters	Name	Type	Range	Default
	<channel>	Channel	CH1|CH2|CH3|CH4|CH5|CH6	–
Usage example	SYST:CHAN:VERS? CH4 / "R2B6"
Related Commands	SYSTem:CHANnel:SLOT? / SYSTem:CHANnel:VERSion?

### SYSTem:COMMunicate:ENABle
Syntax	SYSTem:COMMunicate:ENABle {<bool>}, {<interface>} / SYSTem:COMMunicate:ENABle? {<interface>}
Description	Enables or disables USB, Ethernet, MQTT, NTP service, or the remote service Sockets. The setting is effective after rebooting the BB3. This command setting is not changed by power off or the *RST command. / Enabling the Ethernet interface (SYSTem:COMMunicate:ENABle 1, ETH) will have no effect until a restart is performed using the SYSTem:REStart command. / Remote service Sockets is not supported yet
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	–
	<interface>	Discrete	USB|ETHernet|MQTT|NTP|SOCKets	–
Return	This query returns the status of the selected interface that could be 0 (OFF) or 1 (ON).
Usage example	SYST:COMM:ENAB 0, SER
Related Commands	SYSTem:REStart

### SYSTem:COMMunicate:ETHernet:ADDRess
Syntax	SYSTem:COMMunicate:ETHernet:ADDRess {<ip_address>} / SYSTem:COMMunicate:ETHernet:ADDRess?
Description	Set the IP address of the BB3 manually if the DHCP mode is not enabled. If the DHCP mode is enabled setting the IP address using this command will be ignored.
Parameters	Name	Type	Range	Default
	<ip_address>	Quoted string	15 characters max. formatted as four groups of up to 3 digits (range 0–255), separated by “.”	–
Return	The query returns the current IP address sets manually or assigned by the DHCP server.
Usage example	SYST:COMM:ETH:ADDR? / "192.168.10.100"
Related Commands	SYSTem:COMMunicate:ETHernet:DHCP

### SYSTem:COMMunicate:ETHernet:DHCP
Syntax	SYSTem:COMMunicate:ETHernet:DHCP {<bool>} / SYSTem:COMMunicate:ETHernet:DHCP?
Description	Enable or disable the DHCP mode. In DHCP mode, the DHCP server in the current network assigns network parameters (IP ADdress, DNS address, GATEway address and the Subnet MASk) for the BB3.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	ON
Return	Query the status of the DHCP mode.
Usage example	SYST:COMM:ETH:DHCP? / 1

### SYSTem:COMMunicate:ETHernet:DNS
Syntax	SYSTem:COMMunicate:ETHernet:DNS {<ip_address>} / SYSTem:COMMunicate:ETHernet:DNS?
Description	Set the DNS (Domain Name Service) address if the DHCP mode is not enabled. If the DHCP mode is enabled setting the DNS address using this command will be ignored.
Parameters	Name	Type	Range	Default
	<ip_address>	Quoted string	15 characters max. formatted as four groups of up to 3 digits (range 0–255), separated by “.”	–
Return	The query returns the DNS address sets manually or assigned by the DHCP server. / If DHCP is used query will return “unknown”.
Usage example	SYST:COMM:ETH:DNS "192.168.1.200"
Related Commands	SYSTem:COMMunicate:ETHernet:DHCP

### SYSTem:COMMunicate:ETHernet:GATEway
Syntax	SYSTem:COMMunicate:ETHernet:GATEway {<ip_address>} / SYSTem:COMMunicate:ETHernet:GATEway?
Description	Set the network gateway address if the DHCP mode is not enabled. If the DHCP mode is enabled setting the network gateway address using this command will be ignored.
Parameters	Name	Type	Range	Default
	<ip_address>	Quoted string	15 characters max. formatted as four groups of up to 3 digits (range 0–255), separated by “.”	–
Return	The query returns the network gateway address sets manually or assigned by the DHCP server. If DHCP is used query will return “unknown”.
Usage example	SYST:COMM:ETH:GATE? / "192.168.10.1"
Related Commands	SYSTem:COMMunicate:ETHernet:DHCP

### SYSTem:COMMunicate:ETHernet:HOSTname
Syntax	SYSTem:COMMunicate:ETHernet:HOSTname {<hostname>} / SYSTem:COMMunicate:ETHernet:HOSTname?
Description	This command sets the BB3 local area network (LAN) connection hostname.
Parameters	Name	Type	Range	Default
	<hostname>	Quoted string	1 – 63 characters	EEZ-BB3
Return	The query returns the BB3 local area network (LAN) connection hostname as a quoted string.
Usage example	SYST:COMM:ETH:HOST? / "bb3-test"
Related Commands	SYSTem:COMMunicate:ETHernet:ADDRess / SYSTem:COMMunicate:ETHernet:DHCP

### SYSTem:COMMunicate:ETHernet:MAC
Syntax	SYSTem:COMMunicate:ETHernet:MAC {<mac_address>} / SYSTem:COMMunicate:ETHernet:MAC?
Description	Use this command to set Ethernet communication port MAC address. You can set any combination of six hexadecimal values separated by “-”. The Ethernet connection will work as long as two different machine in the LAN don't have the same MAC address.
Parameters	Name	Type	Range	Default
	<mac_address>	Quoted string	17 characters formatted as six groups of 2 digits separated by “-”	74-69-69-2D-30-00
Return	The query returns the MAC address as a quoted string (six hexadecimal values separated by “-”).
Usage example	SYST:COMM:ETH:MAC? / "70-60-50-40-30-20"

### SYSTem:COMMunicate:ETHernet:PORT
Syntax	SYSTem:COMMunicate:ETHernet:PORT {<number>} / SYSTem:COMMunicate:ETHernet:PORT?
Description	Use this command to change default (5025) Ethernet communication port for SCPI.
Parameters	Name	Type	Range	Default
	<number>	NR1	1-65535	5025
Return	The query returns the Ethernet communication port number.
Usage example	SYST:COMM:ETH:PORT? / 5025
Related Commands	SYSTem:COMMunicate:ENABle / SYSTem:COMMunicate:ETHernet:CONTrol?

### SYSTem:COMMunicate:ETHernet:SMASk
Syntax	SYSTem:COMMunicate:ETHernet:SMASk {<mask>} / SYSTem:COMMunicate:ETHernet:SMASk?
Description	Set the subnet mask if the DHCP mode is not enabled. If the DHCP mode is enabled setting the network gateway address using this command will be ignored.
Parameters	Name	Type	Range	Default
	<mask>	Quoted string	15 characters max. formatted as four groups of up to 3 digits (range 0–255), separated by “.”	–
Return	The query returns the subnet mask sets manually or assigned by the DHCP server. If DHCP is used query will return “unknown”.
Usage example	SYST:COMM:ETH:SMAS "255.255.255.0"
Related Commands	SYSTem:COMMunicate:ETHernet:DHCP

### SYSTem:COMMunicate:MQTT:SETTings
Syntax	SYSTem:COMMunicate:MQTT:SETTings {<address>}, {<port>}, {<user>}, {<password>}, {<period>}
Description	MQTT (MQ Telemetry Transport) is an open OASIS and ISO standard (ISO/IEC PRF 20922) lightweight, publish-subscribe network protocol that transports messages between devices (i.e. clients). / The EEZ BB3 becomes a MQTT client (or node) by establishing connection with this command to the MQTT broker. A broker is a server that receives all messages from the clients and then routes the messages to the appropriate destination clients. It acts as a post office, MQTT doesn’t use the address of the intended recipient but uses the subject line called “Topic”, and anyone who wants a copy of that message will subscribe to that topic. EEZ The BB3 can produce and receive data by both publishing and subscribing. / MQTT connection is initiated by sending a sequence of five comma separated connection parameters.
Parameters	Name	Type	Range	Default
	<address>	Quoted string	Max. 64 characters	–
	<port>	NR1	1 – 99999	–
	<user>	Quoted string	Max. 32 characters	–
	<password>	Quoted string	Max. 32 characters	–
	<period>	NR2	0.1 – 120	1
Usage example	SYST:COMM:MQTT:CONN "farmer.cloudmqtt.com", 1883, "user", "password", 10
Errors	-104,"Data type error"
Related Commands	SYSTem:COMMunicate:ENABle / SYSTem:COMMunicate:MQTT:STATe

### SYSTem:COMMunicate:MQTT:STATe
Syntax	SYSTem:COMMunicate:MQTT:STATe?
Description	This query returns the current status of the MQTT connection.
Return	Returns -1 if an error occurred, 0 if connection is not established, 1 if connection is established. Return codes greater then 1 indicates various transient states.
Usage example	SYST:COMM:MQTT:STAT? / 1
Related Commands	SYSTem:COMMunicate:ENABle / SYSTem:COMMunicate:MQTT:SETTings

### SYSTem:COMMunicate:NTP
Syntax	SYSTem:COMMunicate:NTP {<server>} / SYSTem:COMMunicate:NTP?
Description	Use this command to set the NTP service server network address. The BB3 will try to establish connection with selected NTP service on each power up (hard reset), when *RST is issued or periodically as defined with the SYSTem:COMMunicate:NTP:FREQuency commad.
Parameters	Name	Type	Range	Default
	<server>	Quoted string	IP address or NTP server host name up to 32 characters	europe.pool.ntp.org
Return	The query command returns the NTP service server network address.
Usage example	SYST:COMM:ENAB NTP / SYST:COMM:NTP "ntp.ubuntu.com"
Related Commands	*RST / SYSTem:COMMunicate:ENABle / SYSTem:COMMunicate:FREQuency / SYSTem:DATE / SYSTem:TIME

### SYSTem:COMMunicate:NTP:FREQuency
Syntax	SYSTem:COMMunicate:NTP:FREQuency {<frequency>} / SYSTem:COMMunicate:NTP:FREQuency?
Description	Use this command to set the time in minutes between two connection attempts to the defined NTP service.
Parameters	Name	Type	Range	Default
	<frequency>	NR1	15 – 14400	720
Return	The query command returns the period in minutes between two connection attempts to the NTP service.
Usage example	SYST:COMM:NTP:FREQ 60
Related Commands	*RST / SYSTem:COMMunicate:NTP / SYSTem:DATE / SYSTem:TIME

### SYSTem:COMMunicate:USB:CLAss
Syntax	SYSTem:COMMunicate:USB:CLAss {<usb_class>} / SYSTem:COMMunicate:USB:CLAss?
Description	Use this command to set the USB interface class.
Parameters	Name	Type	Range	Default
	<usb_class>	Discrete	VCOM|MSTOrage|HID	–
Return	The query command returns the USB interface class.
Usage example	SYST:COMM:USB:MODE "DEV" / SYST:COMM:USB:CLA "VCOM"
Related Commands	*RST / SYSTem:COMMunicate:ENABle / SYSTem:COMMunicate:USB:MODE

### SYSTem:COMMunicate:USB:MODE
Syntax	SYSTem:COMMunicate:USB:MODE {<usb_mode>} / SYSTem:COMMunicate:USB:MODE?
Description	Use this command to set the USB interface mode that could be one of the following: / DEVice – the BB3 acts as an USB device when communicate with connected computer. / HOST – in this mode the BB3 provides VBUS power for connected USB device (e.g. mouse, keypad, foot pedal, etc.). The USB class will be automatically set to HID. / OTG – “on-the-go” mode allows detection of USB ID pin (set by USB cable or adapter) and adopt its USB mode and class in accordance with detected device.
Parameters	Name	Type	Range	Default
	<usb_mode>	Discrete	DEVice|HOST|OTG	–
Return	The query command returns the USB interface mode.
Usage example	SYST:COMM:USB:MODE "HOST"
Related Commands	*RST / SYSTem:COMMunicate:ENABle / SYSTem:COMMunicate:USB:CLAss

### SYSTem:COMMunicate:RLSTate
Syntax	SYSTem:COMMunicate:RLSTate {<state>} / SYSTem:COMMunicate:RLSTate?
Description	This command configures the remote/local state of the BB3 according to the following settings: / LOCal – The BB3 is set to front panel and remote interface control. / REMote – The BB3 is set to front panel and remote interface control. / RWLock – The front panel keys are disabled. The BB3 can only be controlled via the remote interface. This programmable setting is completely independent from the front panel lock/unlock function that is available from the front panel menu. If you use this command to lock the front panel, the front panel will be unlocked when AC power is cycled. / The LOCal parameter is the same as SYSTem:LOCal, the REMote parameter is the same as SYSTem:REMote, and the RWLock parameter is the same as SYSTem:RWLock. / The remote/local state is unaffected by *RST or any SCPI commands other than SYSTem:COMMunicate:RLState.
Parameters	Name	Type	Range	Default
	<state>	Discrete	LOCal|REMote|RWLock	LOCal
Return	The query command returns control state of the BB3 that could be LOC, REM, or RWL.
Usage example	SYST:COMM:RLST? / "LOC"
Related Commands	*RST / SYSTem:LOCal / SYSTem:REMote / SYSTem:RWLock

### SYSTem:CPU:FIRMware?
Syntax	SYSTem:CPU:FIRMware?
Description	This query returns BB3 (master) firmware version.
Usage example	SYST:CPU:FIRM? / "1.0"

### SYSTem:CPU:INFOrmation:ONTime:LAST?
Syntax	SYSTem:CPU:INFOrmation:ONTime:LAST?
Description	This query returns time passed after last activation of the BB3. Resolution is 1 minute and this information is stored every 10 minutes in non-volatile memory. Therefore it’s possible that up to 10 minutes is lost after restart caused with power outage or system reset.
Usage example	SYST:CPU:INFO:ONT:LAST? / 15m

### SYSTem:CPU:INFOrmation:ONTime:TOTal?
Syntax	SYSTem:CPU:INFOrmation:ONTime:TOTal?
Description	This query returns total active time of the BB3. Resolution is 1 minute and this information is stored every 10 minutes in non-volatile memory. Therefore it’s possible that up to 10 minutes is lost after restart caused with power outage or system reset.
Usage example	SYST:CPU:INFO:ONT:TOT? / "1h 45m"

### SYSTem:CPU:MODel?
Syntax	SYSTem:CPU:MODel?
Description	This query returns the model name and version of the MCU board. If simulator is used it returns “Simulator”.
Usage example	SYST:CPU:MOD? / "STM32F7, M1 0.3"
Related Commands	SYSTem:CPU:INFOrmation:ONTime:LAST? / SYSTem:CPU:INFOrmation:ONTime:TOTal?

### SYSTem:CPU:SNO?
Syntax	SYSTem:CPU:SNO?
Description	This query returns the BB3 serial number.
Usage example	SYST:CPU:SNO? / "992A003D3338510738323535"
Related Commands	*IDN?

### SYSTem:CPU:VERSion?
Syntax	SYSTem:CPU:VERSion?
Description	This query returns the BB3 version number.
Usage example	SYST:CPU:VERS? / "R2B4"
Related Commands	*IDN?

### SYSTem:DATE
Syntax	SYSTem:DATE {<year>}, {<month>}, {<day>} / SYSTem:DATE?
Description	Sets the date of the system clock (RTC). Specify the year, month, and day. / The self-test procedure compare date and time stored in RTC registers with values stored in the non-volatile memory (EEPROM). When the later is greater then former or any of them lost integrity (i.e. any of value is outside allowed range: for example seconds are higher then 60 or months are higher then 12, etc.) self-test will failed. The *TST? will return 1 and detailed report could be queried using the DIAGnostic:TEST? command. / The bit 3 (TIME) of the Questionable Status register will be set (see Section 3.4) if datetime self-test failed or datetime was never set.
Parameters	Name	Type	Range	Default
	<year>	NR1	2000 – 2099	–
	<month>	NR1	1 – 12	–
	<day>	NR1	1 – 31	–
Return	Query the current date of the system clock in YYYY, MM, DD format.
Usage example	SYST:DATE? / 2015, 10, 24
Related Commands	*TST? / DIAGnostic[:INFOrmation]:TEST? / SYSTem:FORMat:TIME / SYSTem:TIME

### SYSTem:DELay
Syntax	SYSTem:DELay {<time>}
Description	This command inserts pause in parsing and executing SCPI commands. Execution of the next SCPI command will start only when the set delay expires. The programmed value can be up to 10 seconds (set in milliseconds).
Parameters	Name	Type	Range	Default
	<time>	NR1	1 – 10000	–
Usage example	SYST:DEL 1500

### SYSTem:DIGital:INPut:DATA
Syntax	SYSTem:DIGital:INPut:DATA? {<pin>}
Description	This query reads the state of the digital control port. / Applies only to pin 1.
Parameters	Name	Type	Range	Default
	<pin>	NR1	1	1
Return	The query returns the value of the state of input pin.
Usage example	SYST:DIG:INP:DATA? 1 / 1
Related Commands	SYSTem:DIGital:PIN<n>:FUNCtion / SYSTem:DIGital:PIN<n>:POLarity

### SYSTem:DIGital:OUTPut:DATA
Syntax	SYSTem:DIGital:OUTPut:DATA {<pin>}, {<state>} / SYSTem:DIGital:OUTPut:DATA? {<pin>}
Description	This command sets the output data on the digital output pin. Applies only to pin 3 and 4.
Parameters	Name	Type	Range	Default
	<pin>	NR1	3 – 4	–
	<state>	Discrete	ON|OFF|0|1	–
Return	The query returns the last programmed value on the selected pin.
Usage example	SYST:DIG:OUTP:DATA 3, 0
Errors	-114,"Header suffix out of range"
Related Commands	SYSTem:DIGital:PIN<n>:FUNCtion / SYSTem:DIGital:PIN<n>:POLarity

### SYSTem:DIGital:OUTPut:PWM:DUTY
Syntax	SYSTem:DIGital:OUTPut:PWM:DUTY {<pin>}, {<duty>} / SYSTem:DIGital:OUTPut:PWM:DUTY? {<pin>}
Description	Use this command to set duty cycle of the output pin defined as PWM (see the SYSTem:DIGital:PIN<n>:FUNCtion command). Applies only to 4.
Parameters	Name	Type	Range	Default
	<pin>	NR1	4	–
	<duty>	NR3	0 – 100	50
Return	The query returns the square ware generator duty cycle for the specified output pin.
Usage example	SYST:DIG:OUTP:PWM:DUTY 4, 10
Errors	-224,"Illegal parameter value" / -230,"Digital pin function mismatch"
Related Commands	SYSTem:DIGital:OUTPut:PWM:FREQuency / SYSTem:DIGital:PIN<n>:FUNCtion

### SYSTem:DIGital:OUTPut:PWM:FREQuency
Syntax	SYSTem:DIGital:OUTPut:PWM:FREQuency {<pin>}, {<frequency>} / SYSTem:DIGital:OUTPut:PWM:FREQuency? {<pin>}
Description	Use this command to set frequency of the output pin defined as PWM (see the SYSTem:DIGital:PIN<n>:FUNCtion command). Applies only to 4.
Parameters	Name	Type	Range	Default
	<pin>	NR1	4	–
	<frequency>	NR3	0.03 – 5000000	0
Return	The query returns the square ware generator frequency for the specified output pin.
Usage example	SYST:DIG:OUTP:PWM:FREQ 4, 1000
Errors	-224,"Illegal parameter value" / -230,"Digital pin function mismatch"
Related Commands	SYSTem:DIGital:OUTPut:PWM:DUTY / SYSTem:DIGital:PIN<n>:FUNCtion

### SYSTem:DIGital:PIN<n>:FUNCtion
Syntax	SYSTem:DIGital:PIN<n>:FUNCtion {<function>} / SYSTem:DIGital:PIN<n>:FUNCtion?
Description	Use this command to set function of the selected digital port pin. The pin function is saved in non-volatile memory. / All input functions applies only to pin 1 and 2 and all output functions applies only to pin 3 and pin 4. / DINPut – The pin is in digital input mode. / DLOGtrig – The pin is configured as a DLOG trigger input. / DOUTput – The pin is in digital output mode. / FAULt – Setting FAULt means that pin functions as an isolated fault output. The fault signal is true when any output is in a protected state (from OCP, OVP, OTP, OPP) or Fan fault is detected. / INHibit – When pin is configured as an inhibit input, a true signal at the pin will disable all output channels. / ONCouple – output pin synchronize channel output state. / PWM – (pin 4 only) square wave generator output. / SYSTrig – The pin is configured as a system trigger input. When configured as a trigger input, the pin can be selected as the source for trigger signals. See TRIGger[:SEQuence]:SOURce. / TOUTput – This allows a BUS trigger to be sent to any digital port pin that has been configured as a trigger output. A trigger out pulse is generated when the state is on and a bus trigger is received. A BUS trigger is generated using the *TRG command.
Parameters	Name	Type	Range	Default
	<function>	Discrete	DINP|DLOG|DOUT|FAUL|INH|ONC|PWM|SYST|TOUT	–
Return	The query command returns DINP, DOUT, FAUL, INH or TINP.
Usage example	SYST:DIG:PIN1:FUNC INH / SYST:DIG:PIN2:FUNC TINP / SYST:DIG:PIN3:FUNC FAUL / SYST:DIG:PIN4:FUNC ONC
Errors	-114,"Header suffix out of range"
Related Commands	TRIGger[:SEQuence]:SOURce / SYSTem:DIGital:INPut:DATA / SYSTem:DIGital:OUTPut:DATA / SYSTem:DIGital:OUTPut:PWM:DUTY / SYSTem:DIGital:OUTPut:PWM:FREQuency

### SYSTem:DIGital:PIN<n>:POLarity
Syntax	SYSTem:DIGital:PIN<n>:POLarity {<polarity>} / SYSTem:DIGital:PIN<n>:POLarity?
Description	This command sets the polarity of the selected digital port pin. The pin polarity is saved in non-volatile memory. / POSitive – a logical true signal is a voltage high at the pin. For trigger inputs and outputs, POSitive means a rising edge. / NEGative – a logical true signal is a voltage low at the pin. For trigger inputs and outputs, NEGative means a falling edge.
Parameters	Name	Type	Range	Default
	<polarity>	Discrete	POSitive|NEGative	–
Return	The query command returns POS or NEG.
Usage example	SYST:DIG:PIN1:POL POS / SYST:DIG:PIN2:POL NEG
Errors	-114,"Header suffix out of range"
Related Commands	TRIGger[:SEQuence]:SOURce / SYSTem:DIGital:INPut:DATA / SYSTem:DIGital:OUTPut:DATA

### SYSTem:ERRor
Syntax	SYSTem:ERRor[:NEXT]?
Description	This query command reads and clear errors from the error queue. A record of up to 20 errors can be stored in the BB3’s error queue. See also “Error Messages” in Section 7. / Errors are retrieved in first-in-first-out (FIFO) order. The first error returned is the first error that was stored. The BB3 beeps once each time an error is generated. The error queue is cleared when power has been off or after a *CLS command.
Return	SYSTem:ERRor[:NEXT]? queries and clears the error messages in the error queue. The query returns the number and content of the error message.
Usage example	SYST:ERR? / -113,"Undefined header"
Errors	If more than 20 errors have occurred, the last error stored in the queue (the most recent error) is replaced with: / -350,"Queue overflow" / No additional errors are stored until you remove errors from the queue.
Related Commands	*CLS / *RST / SYSTem:ERRor:COUNt

### SYSTem:ERRor:COUNt?
Syntax	SYSTem:ERRor:COUNt?
Description	This query command queries the error/event queue for the number of unread items. As errors and events may occur at any time, more items may be present in the queue at the time it is actually read.
Usage example	SYST:ERR:COUN? / 10
Related Commands	*CLS / *RST / SYSTem:ERRor[:NEXT]

### SYSTem:FAN:SPEed?
Syntax	SYSTem:FAN:SPEed?
Description	Use this query to obtain cooling fan speed in rpm.
Usage example	SYST:FAN:SPE? / 1450
Related Commands	SYSTem:FAN:STATus?

### SYSTem:FAN:STATus?
Syntax	SYSTem:FAN:STATus?
Description	Use this query to obtain information about cooling fan state. / Cooling fan is periodically tested while it’s working (that depends of measured channel’s temperature sensor value). When it does not passed the test, programmed output current is automatically limited to 2 A.
Return	Returns numerical status of the cooling fan: / 0 – fault is detected / 1 – fan is up and running / 2 – fan testing is in progress / 3 – fan is not installed
Usage example	SYST:FAN:STAT? / 1
Related Commands	SYSTem:FAN:SPEed?

### SYSTem:FORMat:DATE
Syntax	SYSTem:FORMat:DATE {<format>} / SYSTem:FORMat:DATE?
Description	Set format that will be used for displaying date.
Parameters	Name	Type	Range	Default
	<format>	Discrete	DMY|MDY	–
Return	The query command returns “DMY” or “MDY”.
Usage example	SYST:FORM:DATE MDY
Related Commands	SYSTem:DATE / SYSTem:FORMat:DATE / SYSTem:FORMat:TIME

### SYSTem:FORMat:TIME
Syntax	SYSTem:FORMat:TIME {<format>} / SYSTem:FORMat:TIME?
Description	Set clock format that will be used for displaying time.
Parameters	Name	Type	Range	Default
	<format>	NR1	12|24	–
Return	The query command returns 12 or 24 clock format.
Usage example	SYST:FORM:TIME 12
Related Commands	SYSTem:DATE / SYSTem:FORMat:DATE / SYSTem:FORMat:TIME

### SYSTem:INHibit?
Syntax	SYSTem:INHibit?
Description	Use this command to query BB3’s inhibit state that is controlled with digital input (see [SOURce]:DIGital:PIN<n>:FUNCtion command). When BB3 is in inhibited state, all channel’s output will be disabled (OFF) and triggered action (e.g. LIST) will be paused. / Channel output state command (OUTPut[:STATe]) can be used when BB3 is in inhibited mode but it will not affect output (i.e. change it to enable state).
Return	Returns 0 if inhibit is not active or 1 if inhibit is active.
Usage example	SYST:INH? / 0
Related Commands	OUTPut[:STATe] {<bool>} [, <channel>] / [SOURce]:DIGital:PIN<n>:FUNCtion

### SYSTem:KLOCk
Syntax	SYSTem:KLOCk
Description	This command similar to the SYSTem:REMote command disables all front-panel options except for [Lock/Unlock] icon. You can push and hold the [Lock/Unlock] for a few seconds to unlock the front panel. The system password may be needed if it is set. / Lock state is saved in non-volatile memory. Therefore, the front panel remains locked even after AC power is cycled.
Usage example	SYST:KLOC
Related Commands	SYSTem:COMMunicate:RLSTate / SYSTem:REMote

### SYSTem:LOCal
Syntax	SYSTem:LOCal
Description	This command places the BB3 in the local mode during remote operation. All options on the front panel are fully functional.
Usage example	SYST:LOC
Related Commands	SYSTem:COMMunicate:RLSTate / SYSTem:REMote / SYSTem:RWLock

MEASure[:SCALar]:TEMPerature[:THERmistor][:DC]
Syntax	MEASure[:SCALar]:TEMPerature[:THERmistor][:DC]? {<sensor>}
Description	Query the temperature measured using the specified temperature sensors as follows: / AUX – auxiliary temperature sensor (built-in fan controller) used to monitor temperature inside BB3 chassis / CH1, CH2, CH3, CH4, CH5, CH6 – temperature sensor connected to channel power module used to regulate fan speed / All installed sensors are periodically tested. When sensor does not passed the test, programmed output current is automatically limited to 2 A. If load that draws more current then 2 A output current will be set to zero.
Parameters	Name	Type	Range	Default
	<sensor>	Discrete	AUX|CH1|CH2|CH3|CH4|CH5|CH6	AUX
Return	Returns the average temperature value in degrees Celsius (oC) as decimal number (NR2). If self-test detect selected temperature sensor failure or sensor is not installed a -240 or -241 error will be generated.
Usage example	MEAS:TEMP? CH2 / 49
Errors	-240,"Hardware error" / -241,"Hardware missing"
Related Commands	DIAGnostic[:INFOrmation]:FAN? / SYSTem:CPU:OPTion? / SYSTem:TEMPerature:PROTection[:HIGH][:LEVel] / SYSTem:TEMPerature:PROTection[:HIGH]:DELay[:TIME] / SYSTem:TEMPerature:PROTection[:HIGH]:STATe / SYSTem:TEMPerature:PROTection[:HIGH]:TRIPped?

### SYSTem:MEASure[:SCALar]:TEMPerature[:THERmistor][:DC]
Syntax	SYSTem:MEASure[:SCALar]:TEMPerature[:THERmistor][:DC]? {<sensor>}
Description	Query the temperature measured using the specified temperature sensors as follows: / AUX – auxiliary temperature sensor (built-in fan controller) used to monitor temperature inside BB3 chassis / CH1, CH2, CH3, CH4, CH5, CH6 – temperature sensor connected to channel power module used to regulate fan speed / All installed sensors are periodically tested. When sensor does not passed the test, programmed output current is automatically limited to 2 A. If load that draws more current then 2 A output current will be set to zero.
Parameters	Name	Type	Range	Default
	<sensor>	Discrete	AUX|CH1|CH2|CH3|CH4|CH5|CH6	AUX
Return	Returns the average temperature value in degrees Celsius (oC) as decimal number (NR2). If self-test detect selected temperature sensor failure or sensor is not installed a -240 or -241 error will be generated.
Usage example	SYST:MEAS:TEMP? CH2 / 49
Errors	-240,"Hardware error" / -241,"Hardware missing"
Related Commands	DIAGnostic[:INFOrmation]:FAN? / SYSTem:CPU:OPTion? / SYSTem:TEMPerature:PROTection[:HIGH][:LEVel] / SYSTem:TEMPerature:PROTection[:HIGH]:DELay[:TIME] / SYSTem:TEMPerature:PROTection[:HIGH]:STATe / SYSTem:TEMPerature:PROTection[:HIGH]:TRIPped?

### SYSTem:MEASure[:SCALar][:VOLTage][:DC]?
Syntax	SYSTem:MEASure[:SCALar][:VOLTage][:DC]? {<device>}
Description	Returns voltage of the RTC (Real-time-clock) backup battery.
Parameters	Name	Type	Range	Default
	<device>	Discrete	RTC	RTC
Usage example	SYST:MEAS? RTC / 3.07

### SYSTem:PASSword:CALibration:RESet
Syntax	SYSTem:PASSword:CALibration:RESet
Description	This command resets the calibration password to the firmware default setting, which is “eezbb3”. This command does not reset the system password.
Usage example	SYST:PASS:CAL:RES
Related Commands	CALibration[:MODE] / CALibration:PASSword:NEW

### SYSTem:PASSword:FPANel:RESet
Syntax	SYSTem:PASSword:FPANel:RESet
Description	This command resets the front panel lockout password to the firmware default setting, which is empty space (“”). This command does not reset the calibration password.
Usage example	SYST:PASS:FPAN:RES
Related Commands	SYSTem:KLOCk / SYSTem:PASSword:NEW

### SYSTem:PASSword:NEW
Syntax	SYSTem:PASSword:NEW {<old>}, {<new>}
Description	Enter a new system password. To change the password, first unsecure the BB3 using the old password. Then, the new code has to be entered. The calibration code may contain up to 16 characters over the remote interface. Minimum length is 4 characters. / The new password is automatically stored in non-volatile memory
Parameters	Name	Type	Range	Default
	<old>	Quoted string	0 to 16 characters	–
	<new>	Quoted string	4 to 16 characters	–
Usage example	SYST:PASS:NEW "","mypass2016"
Errors	122,"Invalid sys password" / 125,"Sys password too long" / 126,"Sys password too short"
Related Commands	SYSTem:KLOCk

### SYSTem:PON:OUTPut:DISable
Syntax	SYSTem:PON:OUTPut:DISable {<bool>} / SYSTem:PON:OUTPut:DISable?
Description	This command controls status off all channel outputs on power up. If enabled (ON), all outputs will be disabled regardless of what is stored in user profile selected for auto recall.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	OFF
Return	Query returns status of forced output disabling on power up.
Usage example	OUTP? / 1 / SYST:PON:OUTP:DIS 1 / (Restart) / OUTP? / 0
Related Commands	MEMory:STATe:RECall:AUTO / SYSTem:POWer

### SYSTem:POWer
Syntax	SYSTem:POWer {<bool>} / SYSTem:POWer?
Description	This command controls powering down and powering up sequence of the AC power inputs. The “Standby” indicator will be switched on when the BB3 enters the standby mode.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	OFF
Return	Query returns BB3 power standby status.
Usage example	SYST:POW ON / SYST:POW? / 1
Related Commands	*RST / *TST? / SYSTem:BEEP:STATe

### SYSTem:POWer:PROTection:TRIP
Syntax	SYSTem:POWer:PROTection:TRIP {<bool>} / SYSTem:POWer:PROTection:TRIP?
Description	Use this command to shut down (set to standby mode) the BB3 when any of channel’s protection tripped. The SYSTem:POWer command has to be used to power the BB3 on again.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	OFF
Return	Query returns status of shutdown when tripped mode.
Usage example	SYST:POW:PROT:TRIP? / 1
Related Commands	SYSTem:POWer

### SYSTem:RELay:CYCLes?
Syntax	SYSTem:RELay:CYCLes? {<chanlist>}
Modules	SMX
Description	This query returns the cycle count on the specified channels. / The SMX46 module channels are mapped in the following way (n is slot number): / (@n03) – Power relay (max. 250 Vac, 8 A per contact) / (@nrc) – Signal relay (1 A max., 30 Vdc / 0.3 A, 125 Vac) in switch matrix consists of 4 rows (r value) and 6 columns (c value). For example, use (@n23) to address relay for switch in row 2 and column 3. / When list of channels separated with comma is specified no space is allowed.
Parameters	Name	Type	Range	Default
	<chanlist>	ChannelList	See channels mapping described above	–
Usage example	Query cycle count of power relay on the SMX46 installed in slot 2: / SYST:REL:CYCL? (@203) / 44 / Query cycle count of multiple signal relays on the SMX46 installed in slot 3: / SYST:REL:CYCL? (@311,312,313) / 13,44,1023
Related Commands	SYSTem:SLOT[:COUNt]? / SYSTem:SLOT:STATe / SYSTem:SLOT:VERSion?

### SYSTem:REMote
Syntax	SYSTem:REMote
Description	Use this command to place the BB3 into remote mode for USB (Virtual COM) or Ethernet remote control. All front-panel options are disabled except for [Lock/Unlock] icon. You can push and hold the [Lock/Unlock] for a few seconds to unlock the front panel. The system password may be needed if it is set.
Usage example	SYST:REM
Related Commands	SYSTem:COMMunicate:RLSTate / SYSTem:LOCal / SYSTem:RWLock

### SYSTem:REStart
Syntax	SYSTem:REStart
Description	Use this command to initiate low-level (hardware) reset. When executed the power up procedure will start and currently active SCPI session will be lost.
Usage example	SYST:RES
Related Commands	*RST

### SYSTem:RWLock
Syntax	SYSTem:RWLock
Description	Places the BB3 in the remote mode for USB (Virtual COM) or Ethernet remote control. This command is the same as SYSTem:REMote, except that all front panel options are disabled, including the [Lock/Unlock] icon.
Usage example	SYST:RWL
Related Commands	SYSTem:COMMunicate:RLSTate / SYSTem:LOCal / SYSTem:REMote

### SYSTem:SLOT[:COUNt]?
Syntax	SYSTem:SLOT[:COUNt]?
Description	This query returns the number of available slots in BB3 enclosure. That number is always 3.
Usage example	SYST:SLOT? / 3
Related Commands	SYSTem:SLOT:MODel? / SYSTem:SLOT:STATe / SYSTem:SLOT:VERSion?

### SYSTem:SLOT:COLor
Syntax	SYSTem:SLOT:COLor {<slot>}, {<color>} / SYSTem:SLOT:COLor? [<slot>]
Modules	SMX PREL
Description	Use this command to set module header and frame color as defined with <color> value. Use 0 to reset module color to the default value. Default module color coresponds to the module slot number. Available colors and its RGB values can be found in the SYSTem:CHANnel:COLor command description. / This command is only applicable to modules that do not have discrete channel numbers (CH1, CH2 …), otherwise it may generate -241,"Hardware missing" execution error.
Parameters	Name	Type	Range	Default
	<slot>	NR1	1 – 3	–
	<color>	NR1	0 – 24	0
Return	Query the status of module on the selected BB3 slot.
Errors	-222,"Data out of range" / -241,"Hardware missing"
Usage example	SYST:SLOT:COL 1,2
Related Commands	SYSTem:CHANnel:COLor

### SYSTem:SLOT:LABel
Syntax	SYSTem:SLOT:LABel {<slot>}, {<label>} / SYSTem:SLOT:LABel? {<slot>}
Description	Use this command to define the module label on selected <slot>.
Parameters	Name	Type	Range	Default
	<slot>	NR1	1 – 3	–
	<label>	Quoted string	1 – 10 characters	–
Return	Returns the label of the module queried as a quoted string.
Errors	-241,"Hardware missing"
Usage example	SYST:SLOT:LAB 1,"Cooler"
Related Commands	SYSTem:CHANnel:LABel / SYSTem:CHANnel:PIN:LABel

### SYSTem:SLOT:MODel?
Syntax	SYSTem:SLOT:MODel? {<slot>}
Description	This query returns the model name of the module installed into specified slot.
Parameters	Name	Type	Range	Default
	<slot>	NR1	1 – 3	–
Usage example	SYST:SLOT:MOD? 2 / "DCP405"
Related Commands	SYSTem:SLOT[:COUNt]? / SYSTem:SLOT:SNO? / SYSTem:SLOT:STATe / SYSTem:SLOT:VERSion?

### SYSTem:SLOT:SNO?
Syntax	SYSTem:SLOT:SNO? {<slot>}
Description	This query returns the serial number of the module installed into specified slot.
Parameters	Name	Type	Range	Default
	<slot>	NR1	1 – 3	–
Usage example	SYST:SLOT:SNO? 1 / "99280045953510720393737"
Related Commands	SYSTem:SLOT:MODel? / SYSTem:SLOT:VERSion?

### SYSTem:SLOT:STATe
Syntax	SYSTem:SLOT:STATe {<slot>}, {<bool>} / SYSTem:SLOT:STATe? [<slot>]
Description	Use this command to enable or disable module installed into specified slot. When disabled an attempt to communicate with module will generate an execution error.
Parameters	Name	Type	Range	Default
	<slot>	NR1	1 – 3	–
	<bool>	Boolean	ON|OFF|0|1	–
Return	Query the status of module on the selected BB3 slot.
Errors	-200,"Execution error"
Usage example	SYST:SLOT:STAT 1,0
Related Commands	*RST / DIAGnostic[:INFOrmation]:TEST? / SYSTem:SLOT[:COUNt]? / SYSTem:SLOT:MODel? / SYSTem:SLOT:VERSion?

### SYSTem:SLOT:VERSion?
Syntax	SYSTem:SLOT:VERSion? {<slot>}
Description	This query returns the version number of the module installed into specified slot.
Parameters	Name	Type	Range	Default
	<slot>	NR1	1 – 3	–
Usage example	SYST:SLOT:VERS? 2 / "R2B11"
Related Commands	*RST / DIAGnostic[:INFOrmation]:TEST? / SYSTem:CHANnel[:COUNt]?

### SYSTem:TEMPerature:PROTection[:HIGH][:LEVel]
Syntax	SYSTem:TEMPerature:PROTection[:HIGH][:LEVel] {<temperature>} [, <sensor>] / SYSTem:TEMPerature:PROTection[:HIGH][:LEVel]? [<sensor>]
Description	Set the over-temperature protection (OTP) value in degrees Celsius (oC) of the selected temperature sensor. When the over-temperature protection function of the specified temperature sensor is enabled (SYSTem:TEMPerature:PROTection[:HIGH]:STATe), one of the following action will be performed when the temperature exceeds the over-temperature protection value currently set: / AUX – Switch off power of the main transformer and set bit 4 of the Questionable Status register / CH1, CH2, CH3, CH4, CH5, CH6 – Disable channel output (OUTPut OFF) and set bit 4 of the Questionable Instrument Isummary register. / If any of above mentioned temperature sensors cause over-temperature condition an error tone will also follow if beeper is enabled (see SYSTem:BEEPer:STATe). / SYSTem:TEMPerature:PROTection[:HIGH]:TRIPped? command can be used to query whether over-temperature protection occurred on the selected temperature sensor.
Parameters	Name	Type	Range	Default
	<temperature>	NR2	10 – 100	70 for CH1 to CH2, and 50 for AUX
	<sensor>	Discrete	AUX|CH1|CH2|CH3|CH4|CH5|CH6	AUX
Return	Query the over-temperature protection (OTP) value of the selected temperature sensor.
Usage example	SYST:TEMP:PROT 50, AUX / SYST:TEMP:PROT? / 50
Related Commands	*RST / SYSTem:TEMPerature:PROTection[:HIGH]:STATe / SYSTem:TEMPerature:PROTection[:HIGH]:TRIPped?

### SYSTem:TEMPerature:PROTection[:HIGH]:CLEar
Syntax	SYSTem:TEMPerature:PROTection[:HIGH]:CLEar [<sensor>]
Description	This command clears the latched protection status when an over-temperature is detected. / All conditions that generate the fault must be removed before the latched status can be cleared. The output is restored to the state it was in before the fault condition occurred.
	Name	Type	Range	Default
	<sensor>	Discrete	AUX|CH1|CH2|CH3|CH4|CH5|CH6	AUX
Usage example	SYST:TEMP:PROT:CLE
Related Commands	SYSTem:TEMPerature:PROTection[:HIGH]:TRIPped

### SYSTem:TEMPerature:PROTection[:HIGH]:DELay[:TIME]
Syntax	SYSTem:TEMPerature:PROTection[:HIGH]:DELay[:TIME] {<delay>} [, <sensor>] / SYSTem:TEMPerature:PROTection[:HIGH]:DELay[:TIME]? [<sensor>]
Description	This command sets the over-temperature protection delay. The over-temperature protection function will not be triggered during the delay time. After the delay time has expired, the over-temperature protection function will be active. / Programmed values can range from 0 to 300 seconds. See also Section 8.1
Parameters	Name	Type	Range	Default
	<delay>	NR1	0 – 300 seconds	10
	<sensor>	Discrete	AUX|CH1|CH2|CH3|CH4|CH5|CH6	AUX
Return	The query returns programmed over-temperature protection delay.
Usage example	SYST:TEMP:PROT:DEL 30, CH2
Related Commands	*RST / SYSTem:TEMPerature:PROTection[:HIGH][:LEVel]

### SYSTem:TEMPerature:PROTection[:HIGH]:STATe
Syntax	SYSTem:TEMPerature:PROTection[:HIGH]:STATe {<bool>} [, <sensor>] / SYSTem:TEMPerature:PROTection[:HIGH]:STATe? [<sensor>]
Description	This command enables or disables the over-temperature protection (OTP) function. The enabled state is ON (1); the disabled state is OFF (0). If the over-temperature protection function is enabled and the measured output power reach value set by [SOURce[<n>]]:POWer:PROTection[:LEVel] the output is disabled and the Questionable Condition status register OPP bit 10 is set.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	OFF
	<sensor>	Discrete	AUX|CH1|CH2|CH3|CH4|CH5|CH6	AUX
Return	The query command returns 0 if the current protection state is OFF, and 1 if the current protection state is ON.
Usage example	SYST:TEMP:PROT:STAT? CH1 / 0
Related Commands	*RST / SYSTem:TEMPerature:PROTection[:HIGH]:CLEar

### SYSTem:TEMPerature:PROTection[:HIGH]:TRIPped?
Syntax	SYSTem:TEMPerature:PROTection[:HIGH]:TRIPped? [<sensor>]
Description	Query whether OTP occurred on the selected temperature sensor. When protection is tripped bit 4 (TEMPerature) of the Questionable Status register will be set (see Section 3.4). / The SYSTem:TEMPerature:PROTection[:HIGH]:CLEar command can be send to clear OTP condition caused by the selected temperature sensor.
Parameters	Name	Type	Range	Default
	<sensor>	Discrete	AUX|CH1|CH2|CH3|CH4|CH5|CH6	AUX
Return	This query returns a 1 if the protection circuit is tripped and a 0 if it is untripped.
Usage example	SYST:TEMP:PROT:TRIP? / 0
Related Commands	SYSTem:TEMPerature:PROTection[:HIGH]:CLEar / SYSTem:TEMPerature:PROTection[:HIGH]:STATe

### SYSTem:TIME
Syntax	SYSTem:TIME {<hours>}, {<minutes>}, {<seconds>} / SYSTem:TIME?
Description	Sets the time of the system clock (RTC). Specify the hours, minutes, and seconds. / The self-test procedure compare date and time stored in RTC registers with values stored in the non-volatile memory (EEPROM). When the later is greater then former or any of them lost integrity (i.e. any of value is outside allowed range: for example seconds are higher then 60 or months are higher then 12, etc.) self-test will failed. The *TST? will return 1 and detailed report could be queried using the DIAGnostic:TEST? command. / The bit 3 (TIME) of the Questionable Status register will be set (see Section 3.4) if datetime self-test failed or datetime was never set.
Parameters	Name	Type	Range	Default
	<hours>	NR1	0 – 23	–
	<minutes>	NR1	0 – 59	–
	<seconds>	NR1	0 – 59	–
Return	Query the current time of the system clock in HH, MM, SS format.
Usage example	SYST:TIME? / 15, 10, 33
Related Commands	*TST? / DIAGnostic[:INFOrmation]:TEST? / SYSTem:DATE / SYSTem:FORMat:TIME

### SYSTem:TIME:DST
Syntax	SYSTem:TIME:DST {<rules>} / SYSTem:TIME:DST?
Description	Use this command to define daylight saving time (DST) rules used in your region. / Firmware v1.0 support limited number of region: Europe, US/Canada and Australia/New Zealand.
Parameters	Name	Type	Range	Default
	<rules>	Discrete	OFF|EU|USA|AUS	OFF
Return	Query the DTS rule name used for adjust time.
Usage example	SYST:TIME:DST EU
Related Commands	SYSTem:COMMunicate:ENABle / SYSTem:COMMunicate:NTP / SYSTem:TIME

### SYSTem:TIME:ZONE
Syntax	SYSTem:TIME:ZONE {<zone>} / SYSTem:TIME:ZONE?
Description	Use this command to define time zone as offset from GMT.
Parameters	Name	Type	Range	Default
	<zone>	Quoted string	-12:00 to 14:00	–
Return	Query the time zone as offset from GMT in the following format: “[sign]hh:mm GMT”
Usage example	SYST:TIME:ZONE 1 / SYST:TIME:ZONE? / "+01:00 GMT"
Related Commands	SYSTem:COMMunicate:ENABle / SYSTem:COMMunicate:NTP / SYSTem:TIME

### SYSTem:VERSion?
Syntax	SYSTem:VERSion?
Description	This command returns the version of the SCPI (Standard Commands for Programmable Instruments) standard with which the instrument is in compliance
Return	The command returns a string in the form “YYYY.V”, where YYYY represents the year of the version and V represents a version for that year.
Usage example	SYST:VERS? / 1999.0