## SENSe
The SENSe control the current measurement range, energy counting/window, the data acquisition sequence and DLOG viewer parameters.

SCPI command	Description
SENSe
:CURRent
[:DC]
RANGe {<range>} [, <channel>]	Selects a DC current measurement range
:DLOG
:CLEar	Resets DLOG settings
:FUNCtion
:CURRent {<bool>}, {<channel>}	Enables/disables output current internal data logging
:POWer {<bool>}, {<channel>}	Enables/disables output power internal data logging
:VOLTage {<bool>}, {<channel>}	Enables/disables output voltage internal data logging
:PERiod {<time>}	Sets the sample period for internal data logging
:TIME {<time>}	Sets the sample duration for internal data logging
:TRACe
[:DATA] {<value>}[, ...]	Adds new data into log file
:REMark {<user remark>}	Adds user description to the log file
:X
:LABel {<label>}	Sets X-axis label
[:RANGe]:MIN {<min>}	Sets X-axis min. value
[:RANGe]:MAX {<max>}	Sets X-axis max. value
SCALe
:STEP {<step>}	Sets X-axis step value
:UNIT {<unit>}	Sets X-axis units
:Y<n>
:LABel {<label>}	Sets Y-axis<n> label
[:RANGe]:MIN {<min>}	Sets Y-axis<n> min. value
[:RANGe]:MAX {<max>}	Sets Y-axis<n> max. value
:SCALe
:UNIT {<unit>}	Sets Y-axis<n> units

### SENSe:CURRent[:DC]:RANGe
Syntax	SENSe:CURRent[:DC]:RANGe {<range>} [, <channel>] / SENSe:CURRent[:DC]:RANGe? [<channel>]
Modules	DCP
Description	This command selects a DC current measurement range when Power board with multiple current ranges is installed (e.g. DCP405 that can be find out with the SYSTem:CHANnel:MODel? command). The entered value must be higher than the maximum current that you expect to measure. Units are in amperes.
Parameters	Name	Type	Range	Default
	<range>	NR2|Discrete	0.5, 5, LOW|HIGH|BEST	BEST
	<channel>	Channel	CH1|CH2|CH3|CH4|CH5|CH6	–
Usage example	SENS:CURR:RANG? / 0.05
Errors	-241,"Hardware missing"
Related Commands	SYSTem:CHANnel:MODel?

### SENSe:DLOG:CLEar
Syntax	SENSe:DLOG:CLEar
Description	This command resets all DLOG parameters.
Usage example	SENS:DLOG:CLE
Errors	-241,"Hardware missing"

### SENSe:DLOG:FUNCtion:CURRent
Syntax	SENSe:DLOG:FUNCtion:CURRent {<bool>}, {<channel>} / SENSe:DLOG:FUNCtion:CURRent? {<channel>}
Description	Use this command to enable or disable output current internal data logging on the specified channel.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	OFF
	<channel>	Channel	CH1|CH2|CH3|CH4|CH5|CH6	–
Return	The query command returns the status (0 or 1) of the output current internal data logging enabled on the specified channel
Usage example	SENS:DLOG:FUNC:CURR ON, CH1
Related Commands	INIT:DLOG / SENSe:DLOG:FUNCtion:POWer / SENSe:DLOG:FUNCtion:VOLTage / SENSe:DLOG:PERiod / SENSe:DLOG:TIME

### SENSe:DLOG:FUNCtion:POWer
Syntax	SENSe:DLOG:FUNCtion:POWer {<bool>}, {<channel>} / SENSe:DLOG:FUNCtion:POWer? {<channel>}
Description	Use this command to enable or disable output power internal data logging on the specified channel.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	OFF
	<channel>	Channel	CH1|CH2|CH3|CH4|CH5|CH6	–
Return	The query command returns the status (0 or 1) of the output power internal data logging enabled on the specified channel
Usage example	SENS:DLOG:FUNC:VOLT OFF, CH2
Related Commands	INIT:DLOG / SENSe:DLOG:FUNCtion:CURRent / SENSe:DLOG:FUNCtion:VOLTage / SENSe:DLOG:PERiod / SENSe:DLOG:TIME

### SENSe:DLOG:FUNCtion:VOLTage
Syntax	SENSe:DLOG:FUNCtion:VOLTage {<bool>}, {<channel>} / SENSe:DLOG:FUNCtion:VOLTage? {<channel>}
Description	Use this command to enable or disable output voltage internal data logging on the specified channel.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	OFF
	<channel>	Channel	CH1|CH2|CH3|CH4|CH5|CH6	–
Return	The query command returns the status (0 or 1) of the output voltage internal data logging enabled on the specified channel
Usage example	SENS:DLOG:FUNC:VOLT OFF, CH2
Related Commands	INIT:DLOG / SENSe:DLOG:FUNCtion:CURRent / SENSe:DLOG:FUNCtion:POWer / SENSe:DLOG:PERiod / SENSe:DLOG:TIME

### SENSe:DLOG:PERiod
Syntax	SENSe:DLOG:PERiod {<time>} / SENSe:DLOG:PERiod?
Description	Use this command to set the data logging sample period in seconds.
Parameters	Name	Type	Range	Default
	<time>	NR2	0.005 – 120	0.02
Return	The query command returns the sample period in seconds.
Usage example	SENS:DLOG:PER 1
Related Commands	INITiate:DLOG / SENSe:DLOG:TIME

### SENSe:DLOG:TIME
Syntax	SENSe:DLOG:TIME {<time>} / SENSe:DLOG:TIME?
Description	This command sets the data logging sample duration, the entered value is in seconds. For sample duration, the maximum is 86 400 000 seconds or 24 000 hours or 1000 days (depending of the available space on the installed SD Card, see the MMEMory:INFOrmation? command).
Parameters	Name	Type	Range	Default
	<time>	NR1	1 – 86400000	60
Return	The query command returns the sample duration in seconds.
Usage example	SENS:DLOG:TIME 3600
Related Commands	INITiate:DLOG / MMEMory:INFOrmation? / SENSe:DLOG:PERiod

### SENSe:DLOG:TRACe[:DATA]
Syntax	SENSe:DLOG:TRACe[:DATA] {<value>}[, ...]
Description	When the X-axis and all Y-axes is defined, and when the DLOG file name is specified, it remains to add points (data) into DLOG. Use this command to do this as many times as needed. The first time this command is called, the Y-axis values are set for the first point (as many <value> parameters as the Y-axis should be set), the second time for the second point, etc. / When the values of all points are defined, execute the ABORt:DLOG command and the creation of the DLOG file is complete.
Parameters	Name	Type	Range	Default
	<value>	NR2	Y:RANGe:MINimum to Y:RANGe:MAXimum	–
Usage example	SENS:DLOG:TRAC 10.2, 33.34, 0.022
Related Commands	ABORt:DLOG

### SENSe:DLOG:TRACe:X:LABel
Syntax	SENSe:DLOG:TRACe:X:LABel {<label>} / SENSe:DLOG:TRACe:X:LABel?
Description	Use this command to set DLOG X-axis label name.
Parameters	Name	Type	Range	Default
	<label>	Quoted string	0 to 32 characters	–
Return	The query returns X-axis label name as quoted string.
Usage example	SENS:DLOG:TRAC:X:LAB "Imon"
Related Commands	SENSe:DLOG:TRACe:X:STEP / SENSe:DLOG:TRACe:X:UNIT

### SENSe:DLOG:TRACe:REMark
Syntax	SENSe:DLOG:TRACe:REMark {<user remark>} / SENSe:DLOG:TRACe:REMark?
Description	Use this command to add description to the log file.
Parameters	Name	Type	Range	Default
	<user remark>	Quoted string	0 to 128 characters	–
Return	The query command returns log file description as a quoted string.
Usage example	SENS:DLOG:TRAC:REM "Test data log"
Related Commands	INITiate:DLOG / INITiate:DLOG:TRACe / SENSe:DLOG:TRACe[:DATA]

### SENSe:DLOG:TRACe:X:UNIT
Syntax	SENSe:DLOG:TRACe:X:UNIT {<unit>} / SENSe:DLOG:TRACe:X:UNIT?
Description	This command specifies DLOG X-axis unit, typically this is the time, or SECond.
Parameters	Name	Type	Range	Default
	<unit>	Discrete	VOLT|AMPEr|WATT|JOULe|SECOnd|OHM|FARAd|HERTz	–
Return	The query command returns log file X-axis units name as quoted string.
Usage example	SENS:DLOG:TRAC:X:UNIT VOLT
Related Commands	SENSe:DLOG:TRACe:X:STEP / SENSe:DLOG:TRACe:X:LABel

### SENSe:DLOG:TRACe:X:STEP
Syntax	SENSe:DLOG:TRACe:X:STEP {<step>} / SENSe:DLOG:TRACe:X:STEP?
Description	Use this command to set step or ∆x, i.e. the distance between two points on the X-axis of the DLOG.
Parameters	Name	Type	Range	Default
	<step>	NR2	Greater then zero	–
Return	The query command returns X-axis step value.
Usage example	SENS:DLOG:TRAC:X:STEP 0.1
Related Commands	SENSe:DLOG:TRACe:X:LABel / SENSe:DLOG:TRACe:X:UNIT

### SENSe:DLOG:TRACe:X[:RANGe]:MIN
Syntax	SENSe:DLOG:TRACe:X[:RANGe]:MIN {<min>} / SENSe:DLOG:TRACe:X[:RANGe]:MIN?
Description	Use this command to set the first point of the DLOG X-axis. The next point has the value MIN + STEP, then MIN + 2 * STEP, etc.
Parameters	Name	Type	Range	Default
	<min>	NR2	–	–
Return	The query command returns set X-axis minimal value.
Usage example	SENS:DLOG:TRAC:X:MIN 0
Related Commands	SENSe:DLOG:TRACe:X:UNIT

### SENSe:DLOG:TRACe:X[:RANGe]:MAX
Syntax	SENSe:DLOG:TRACe:X[:RANGe]:MAX {<max>} / SENSe:DLOG:TRACe:X[:RANGe]:MAX?
Description	Optional command to set the value of the last point on the DLOG X-axis.
Parameters	Name	Type	Range	Default
	<max>	NR2	–	–
Return	The query command returns set X-axis maximum value.
Usage example	SENS:DLOG:TRAC:X:MAX 20
Related Commands	SENSe:DLOG:TRACe:X:UNIT

### SENSe:DLOG:TRACe:X:SCALe
Syntax	SENSe:DLOG:TRACe:X:SCALe {<scale>} / SENSe:DLOG:TRACe:X:SCALe?
Description	Use this command to set the X-axis scale. / Currently, only linear (LINear) scale is supported.
Parameters	Name	Type	Range	Default
	<scale>	Discrete	LINear|LOGarithmic	–
Return	The query command returns X-axis scale type.
Usage example	SENS:DLOG:TRAC:X:SCAL LIN
Related Commands	SENSe:DLOG:TRACe:Y:SCALe

### SENSe:DLOG:TRACe:Y<n>:LABel
Syntax	SENSe:DLOG:TRACe:Y<n>:LABel {<label>} / SENSe:DLOG:TRACe:Y<n>:LABel?
Description	Use this command to set DLOG Y-axis label name. DLOG file could contains up to 18 Y-axis defined by <n>.
Parameters	Name	Type	Range	Default
	<label>	Quoted string	0 to 32 characters	–
Return	The query returns label name as quoted string of the selected Y-axis.
Usage example	SENS:DLOG:TRAC:Y1:LAB "Umon"
Related Commands	SENSe:DLOG:TRACe:Y<n>:UNIT

### SENSe:DLOG:TRACe:Y<n>[:RANGe]:MIN
Syntax	SENSe:DLOG:TRACe:Y<n>[:RANGe]:MIN {<min>} / SENSe:DLOG:TRACe:Y<n>[:RANGe]:MIN?
Description	Use this command to set the first point of the <n> Y-axis. The next point has the value MIN + STEP, then MIN + 2 * STEP, etc.
Parameters	Name	Type	Range	Default
	<min>	NR2	–	–
Return	The query command returns set <n> Y-axis minimal value.
Usage example	SENS:DLOG:TRAC:Y1:MIN 1
Related Commands	SENSe:DLOG:TRACe:Y<n>[:RANGe]:MAX / SENSe:DLOG:TRACe:Y<n>:UNIT

### SENSe:DLOG:TRACe:Y<n>[:RANGe]:MAX
Syntax	SENSe:DLOG:TRACe:Y<n>[:RANGe]:MAX {<max>} / SENSe:DLOG:TRACe:Y<n>[:RANGe]:MAX?
Description	Optional command to set the value of the last point on the <n> Y-axis.
Parameters	Name	Type	Range	Default
	<max>	NR2	–	–
Return	The query command returns set <n> Y-axis maximum value.
Usage example	SENS:DLOG:TRAC:Y1:MAX 50
Related Commands	SENSe:DLOG:TRACe:Y<n>[:RANGe]:MIN / SENSe:DLOG:TRACe:Y<n>:UNIT

### SENSe:DLOG:TRACe:Y:SCALe
Syntax	SENSe:DLOG:TRACe:Y:SCALe {<scale>} / SENSe:DLOG:TRACe:Y:SCALe?
Description	Use this command to set the Y-axis scale.
Parameters	Name	Type	Range	Default
	<scale>	Discrete	LINear|LOGarithmic	–
Return	The query command returns Y-axis scale type.
Usage example	SENS:DLOG:TRAC:Y:SCAL LIN
Related Commands	SENSe:DLOG:TRACe:X:SCALe

### SENSe:DLOG:TRACe:Y<n>:UNIT
Syntax	SENSe:DLOG:TRACe:Y<n>:UNIT {<unit>} / SENSe:DLOG:TRACe:Y<n>:UNIT?
Description	This command specifies DLOG <n> Y-axis unit.
Parameters	Name	Type	Range	Default
	<unit>	Discrete	VOLT|AMPEr|WATT|JOULe|SECOnd|OHM|FARAd|HERTz	–
Return	The query command returns <n> Y-axis units name as quoted string.
Usage example	SENS:DLOG:TRAC:Y1:UNIT "AMPE"
Related Commands	SENSe:DLOG:TRACe:Y<n>:LABel / SENSe:DLOG:TRACe:Y<n>:UNIT