## CALibration
This subsystem provides commands for the module output calibration. Only one channel can be calibrated at a time. If calibration mode has not been enabled with CALibration:STATe, the calibration commands will generate an error. Use CALibration:SAVE to save any changes, otherwise all changes will be lost on exit from calibration mode. Within the same calibration session both output voltage and current can be calibrated for the currently selected channel.
Calibration cannot start if channel output is not enabled (OUTPut[:STATe] ON).
During calibration process up to 20 measurement points can be entered within range for selected module and value (e.g. 0 to 40 V for DCP405 power module). The minimum recommended number of measurement points is two to compensate slope and offset errors.

SCPI Command	Description
CALibration[:MODE] {<bool>}, {<password>}	Enables/disables calibration mode
:CLEar {<password>}	Clears all calibration parameters
:CURRent
[:DATA] {<value>}	Enters the calibration value for set point
:LEVel {<point>}, {<level>}	Calibrates the output current programming
:RANGe {range}	Sets current range for multiple current range module
:PASSword
:NEW {<old>}, {<new>}	Changes calibration password
:REMark {<string>}	Saves calibration information
:SAVE	Saves the new cal constants in non-volatile memory
:SCReen:INIT	Initiates touchscreen calibration procedure
:STATe {<bool>}, {<password>}	Enables calibration parameters
:VOLTage
[:DATA] {<value>}	Enters the calibration value for set point
:LEVel {<point>}, {<level>}	Calibrates the output voltage programming

### CALibration[:MODE]
Syntax	CALibration[:MODE] {<bool>}, {<password>} / CALibration[:MODE]?
Modules	DCP DCM SMX
Description	This command enables or disables calibration mode. Calibration mode must be enabled for the channel to accept any calibration commands. The first parameter specifies the ON (1) or OFF (0) state. The second parameter is the password. Successful execution of this command set both output VOLTage and CURRent of the selected channel to the MINimum value (see Section 8.1). / Execution of this command also affects bit 0 (CALibrating) of the Operation Instrument Isummary register (see Section 3.3.2). / If both voltage and current calibration parameters exists on calibration mode exit (CALibration[:MODE] OFF) the CALibration:STATe ON command will automatically follows.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	OFF
	<password>	Quoted string	4 to 16 characters	"eezbb3"
Return	The returned parameter is 0 (OFF) or 1 (ON).
Usage example	See Section 10.2
Errors	102,"Invalid cal password" / 104,"Bad sequence of calibration commands" / 312,"Cannot execute when the channels are coupled"
Related Commands	CALibration:STATe / DIAGnostic[:INFOrmation]:OTIMe? / INSTrument:COUPle:TRACking

### CALibration:CLEar
Syntax	CALibration:CLEar {<password>}
Modules	DCP DCM SMX
Description	Clear all calibration parameters stored in the module non-volatile memory for the currently selected channel. After successful execution of this command CALibration:STATe will be set to OFF (0) and further usage of the calibration data will be disabled. This command will be also filled calibration remark with the date and note that calibration data has been cleared.
Parameters	Name	Type	Range	Default
	<password>	Quoted string	4 to 16 characters	"eezbb3"
Return	None
Usage example	DIAG:CAL? / "remark=2020-04-28 new cal", / "u_cal_params_exists=1", / "u_point1_dac=0.150000", / "u_point1_data=0.145000", / "u_point1_adc=0.178900", / "u_point2_dac=38.000000", / "u_point2_data=39.292000", / "u_point2_adc=38.032799", / "i_5A_cal_params_exists=1", / "i_5A_point1_dac=0.050000", / "i_5A_point1_data=0.060100", / "i_5A_point1_adc=0.059840", / "i_5A_point2_dac=4.800000", / "i_5A_point2_data=5.072900", / "i_5A_point2_adc=4.810040", / "i_50mA_cal_params_exists=1", / "i_50mA_point1_dac=0.000500", / "i_50mA_point1_data=0.000591", / "i_50mA_point1_adc=0.000600", / "i_50mA_point2_dac=0.048000", / "i_50mA_point2_data=0.049897", / "i_50mA_point2_adc=0.048100" / CAL:STAT? / 1 / CAL:CLE "eezbb3" / CAL:STAT? / 0 / DIAG:CAL? / "remark= Not calibrated", / "u_cal_params_exists=0", / "i_cal_params_exists=0"
Errors	102,"Invalid cal password"
Related Commands	CALibration:STATe / DIAGnostic[:INFOrmation]:CALibration?

### CALibration:CURRent[:DATA]
Syntax	CALibration:CURRent[:DATA] {<value>}
Modules	DCP DCM
Description	This command can only be used when calibration is enabled and the output state of the currently selected channel is ON. It enters a current value that is obtained by reading an external meter. The calibration level (CALibration:CURRent:LEVel {<point>, <level>}) has to be selected first for the value being entered. The BB3 then computes new current calibration constants. These constants has to be stored in non-volatile memory with CALibration:SAVE command.
Parameters	Name	Type	Range	Default
	<value>	NR2	-0.5 A to MAX +0.5 A / The maximum value is dependent on the power module current rating. See Section 8.1	–
Return	None
Usage example	See Section 10.2
Errors	104,"Bad sequence of calibration commands" / 107,"Cal value out of range"
Related Commands	CALibration:CURRent:LEVel

### CALibration:CURRent:LEVel
Syntax	CALibration:CURRent:LEVel {<point>}, {<level>}
Modules	DCP DCM
Description	This command can only be used when calibration is enabled and the output state of the currently selected channel is ON. It sets the power module to a calibration point that is entered with the CAL:CURR[:DATA] command.
Parameters	Name	Type	Range	Default
	<point>	NR1	1 – 20	–
	<level>	NR2	0 to MAX (see also Section 8.1)	–
Return	None
Usage example	See Section 10.2
Errors	101,"Calibration state is off" / 104,"Bad sequence of calibration commands"
Related Commands	CALibration:CURRent[:DATA] / CALibration:STATe / INSTrument:NSELect / INSTrument[:SELect] / [SOURce[<n>]]:CURRent

### CALibration:CURRent:RANGe
Syntax	CALibration:CURRent:RANGe {<range>}
Modules	DCP
Description	When BB3 is equipped with power modules that has multiple current range (e.g. DCP405 that can be find out with the SYSTem:CHANnel:MODel? command) it’s recommended to perform calibration of all ranges. Use this command to select current range on which calibration will be accomplished.
Parameters	Name	Type	Range	Default
	<range>	Discrete|NR2	LOW|HIGH|0.05|5	HIGH|5
Return	None
Usage example	See Section 10.2
Errors	101,"Calibration state is off" / -241,"Hardware missing"
Related Commands	CALibration:CURRent:LEVel / CALibration:STATe / SYSTem:CHANnel:MODel?

### CALibration:PASSword:NEW
Syntax	CALibration:PASSword:NEW {<old>}, {<new>}
Modules	DCP DCM SMX
Description	Enter a new calibration password. To change the password, first unsecure the BB3 using the old password. Then, the new code has to be entered. The calibration code may contain up to 16 characters over the remote interface. Minimum length is 4 characters. / The new password is automatically stored in non-volatile memory and does not have to be stored with CALibration:SAVE.
Parameters	Name	Type	Range	Default
	<old>	Quoted string	4 to 16 characters	eezbb3
	<new>	Quoted string	4 to 16 characters	–
Return	None
Usage example	CAL:PASS:NEW "eezbb3", "mycal1234"
Errors	102,"Invalid cal password" / 105,"Cal password too long" / 106,"Cal password too short"

### CALibration:REMark
Syntax	CALibration:REMark {<user remark>} / CALibration:REMark?
Modules	DCP DCM SMX
Description	Record calibration information about the power module. The BB3 should be in calibration mode before sending a calibration message.
Parameters	Name	Type	Range	Default
	<user remark>	Quoted string	0 to 32 characters	"Calibration passed"
Return	Query the calibration message.
Usage example	CAL:REM "Calibrated by EEZ" / CAL:REM? / "Calibrated by EEZ" / See also Section 10.2
Errors	The following errors could be generated by command but not query: / 101,"Calibration state is off" / 104,"Bad sequence of calibration commands"
Related Commands	CALibration:STATe

### CALibration:SAVE
Syntax	CALibration:SAVE
Modules	DCP DCM SMX
Description	This command saves calibration constants in non-volatile memory after the calibration procedure has been completed. If calibration mode is exited by programming CALibration:STATe OFF without first saving the new constants, the previous constants are restored. Execution of this command also affects bit 0 (CALibrating) of the Operation Instrument Isummary register (see Section 3.3.2).
Return	None
Usage example	See Section 10.2
Errors	-340,"Calibration failed" / 101,"Calibration state is off" / 104,"Bad sequence of calibration commands" / 111,"No new cal data exists"

### CALibration:SCReen:INIT
Syntax	CALibration:SCReen:INIT
Description	Use this command to initiate touchscreen calibration procedure when calibration data are lost or corrupted. Calibration has to be performed locally on the BB3. / New calibration procedure can be initiated also by touching the screen and hold for more then 15 seconds.
Return	None
Usage example	CAL:SCR:INIT

### CALibration:STATe
Syntax	CALibration:STATe {<bool>} / CALibration:STATe?
Modules	DCP DCM SMX
Description	This command enables or disables usage of calibration parameters if they exists.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	OFF|ON|0|1	ON
Return	The returned parameter is 0 (OFF) or 1 (ON).
Usage example	CAL:STAT OFF
Errors	110,"Cal params missing or corrupted"
Related Commands	DIAGnostic[:INFOrmation]:OTIMe?

### CALibration:VOLTage[:DATA]
Syntax	CALibration:VOLTage[:DATA] {<value>}
Modules	DCP DCM SMX
Description	This command can only be used when calibration is enabled and the output state of the currently selected channel is ON. It enters a voltage value that you obtained by reading an external meter. The calibration level (see CALibration:VOLTage:LEVel {<point>, <level>}) for the value being entered. The BB3 then computes new voltage calibration constants. These constants has to be stored in non-volatile memory with CALibration:SAVE command.
Parameters	Name	Type	Range	Default
	<value>	NR2	-1 V to MAX +1 V / The maximum value is dependent on the power module voltage rating. See Section 8.1	–
Return	None
Usage example	See Section 10.2
Errors	104,"Bad sequence of calibration commands" / 107,"Cal value out of range"
Related Commands	CALibration:SAVE / CALibration:STATe / CALibration:VOLTage:LEVel / INSTrument:NSELect / INSTrument[:SELect]

### CALibration:VOLTage:LEVel
Syntax	CALibration:VOLTage:LEVel {<point>}, {<level>}
Modules	DCP DCM SMX
Description	This command can only be used when calibration is enabled and the output state of the currently selected channel is ON. It sets the power module to a calibration point that is entered with the CAL:VOLT[:DATA] command.
Parameters	Name	Type	Range	Default
	<point>	NR1	1 – 20	–
	<level>	NR2	0 to MAX (see also Section 8.1)	–
Return	None
Usage example	See Section Section 10.2
Errors	101,"Calibration state is off" / 104,"Bad sequence of calibration commands"
Related Commands	CALibration:STATe / CALibration:VOLTage[:DATA] / INSTrument:NSELect / INSTrument[:SELect] / [SOURce[<n>]]:VOLTage