## MEASure
Measure commands return back the output voltage, current, power or temperature. They trigger the acquisition of new data before returning the reading. Measurements are performed by digitizing the instantaneous output voltage, current or temperature. Output power is calculated as product of measured voltage and current. Keyword [:DC] is optional since all measurement are by default of the DC level of the signal.

SCPI command	Description
MEASure
:DIGital[:BYTE]? [<channel>]	Takes a measurement; returns digital logic state
[:SCALar]
:CURRent
[:DC]? [<channel>]	Takes a measurement; returns the average current
:POWer
[:DC]? [<channel>]	Takes a measurement; returns the average power
[:VOLTage]
[:DC]? [<channel>]	Takes a measurement; returns the average voltage
:TEMPerature
[:THERmistor][:DC]? {<chanlist>}	Takes a measurement; returns the average temperature

### MEASure:DIGital[:BYTE]
Syntax	MEASure:DIGital[:BYTE]? [<channel>]
Modules	MIO
Description	Query the digital logic state of the all 8 digital inputs on the MIO168 module.
Parameters	Name	Type	Range	Default
	<channel>	Channel	(@n01) where n is the slot number	–
Return	Returns a binary-weighted decimal representing the bits set on the selected digital input channel.
Usage example	MEAS:DIG? (@101) / 3
Related Commands	SOURce:DIGital:SPEed / SOURce:DIGital:RANGe

### MEASure[:SCALar]:CURRent[:DC]
Syntax	MEASure[:SCALar]:CURRent[:DC]? [<channel>]
Modules	DCP DCM MIO
Description	Query the current measured across the current sense resistor inside the power module.
Parameters	Name	Type	Range	Default
	<channel>	Channel	CH1|CH2|CH3|CH4|CH5|CH6	–
Return	Returns the average output current in amperes as decimal number (NR2).
Usage example	Measure current on the currently selected channel (CH1) and CH2: / MEAS:CURR?;:MEAS:CURR? CH2 / 1.23;0.12
Related Commands	INSTrument:COUPle:TRACking

### MEASure[:SCALar]:POWer[:DC]
Syntax	MEASure[:SCALar]:POWer[:DC]? [<channel>]
Modules	DCP DCM
Description	Query the output power calculated as product of measured voltage and current value.
Parameters	Name	Type	Range	Default
	<channel>	Channel	CH1|CH2|CH3|CH4|CH5|CH6	–
Return	Returns the average output power in watts as decimal number (NR2).
Usage example	MEAS:POW? CH2 / 80.44
Related Commands

### MEASure[:SCALar][:VOLTage][:DC]
Syntax	MEASure[:SCALar][:VOLTage][:DC]? [<channel>]
Modules	DCP DCM MIO
Description	Query the voltage measured at the sense terminals of the selected channel.
Parameters	Name	Type	Range	Default
	<channel>	Channel	CH1|CH2|CH3|CH4|CH5|CH6	–
Return	Returns the average output voltage in volts as decimal number (NR2).
Usage example	MEAS:VOLT? CH1 / 43.25 / Query voltage of the channel 2 that is currently selected: / INST CH2 / MEAS? / 12.40
Related Commands	INSTrument:COUPle:TRACking

### MEASure[:SCALar]:TEMPerature[:THERmistor][:DC]
Syntax	MEASure[:SCALar]:TEMPerature[:THERmistor][:DC]? [<chanlist>]
Modules	MUX
Description	Query the temperature measured on the temperature sensor of the selected channel.
Parameters	Name	Type	Range	Default
	<chanlist>	ChannelList	MUX: (@n31) where n is the slot number	–
Return	Returns the average output voltage in volts as decimal number (NR2).
Usage example	MEAS:TEMP? (@131) / 29
Related Commands	SYSTem:MEASure[:SCALar]:TEMPerature[:THERmistor][:DC]