# Programming examples
## Set channel output values and working with the OCP
This is a SCPI commands sequence that sets a voltage, current, and the over-current protection (OCP) on the channel two:

INST?	Check currently selected output
1
INST CH2	Select channel two as current channel
VOLT 10	Set output voltage
CURR	Set output current
CURR:PROT:STAT?	Check OCP status
0
CURR:PROT:STAT 1	Enable OCP
CURR:PROT:DEL 100ms	Set OCP delay
OUTP 1	Enable output
MEAS?	Measure output voltage
10.00
MEAS:CURR?	Measure output current
0.00	Current is zero since no load is connected

If software simulator is used, connection of the load can be also simulated:

SIMU:LOAD 20	Define connected load impedance
MEAS?	Measure voltage once again
10.00
MEAS:CURR?	Measure current once again
0.50	Measured current

The following command sequence could be used to test channel mode with load previously defined and after the load impedance is lowered enough that output current reach programmed value. The OCP has to be disabled because previously defined 100 ms delay does not give us enough time to execute the whole sequence for testing channel mode and output voltage and current values:

OUTP:MODE?	Check mode of operation
"CV"	The channel is in constant-voltage mode since output current is below previously programmed level
SIMU:LOAD?	Check load value
10
CURR:PROT:STAT?	Check OCP status
1
CURR:PROT:STAT OFF	Disable OCP
SIMU:LOAD 4	Decrease load impedance
OUTP:MODE?	Check once again mode of operation
"CC"	Channel enters constant-current mode since Imax = U / R = 10 / 4 = 2.5 A and current is limited to the 1 A
MEAS:CURR?	Measure output current
1.00
MEAS?	Measure output voltage
4.00	Output voltage is decreased since U = I * R = 1 * 4 = 4 V

The OCP will “trip” when output current reach programmed value and channel stay in the CC mode for more then programmed OCP delay time. To test that with e.g. the software simulator we’ll disable channel output first, enable OCP and when change channel output back to enabled state:

OUTP OFF	Disable channel output
CURR:PROT:TRIP?	Check OCP status
0	OCP is not activated
CURR:PROT:STAT ON	Enable OCP
VOLT?	Check programmed output voltage
10.00
CURR?	Check programmed output current
1.00
SIMU:LOAD?	Check simulated load value
4
OUTP ON	Enable channel output
CURR:PROT:TRIP?	Check OCP status once again
1	OCP has been tripped
OUTP?	Check channel output state
0	Channel output is changed to OFF stated by the OCP

The channel output state cannot be changed to enabled until any of protection is active. We have to clear protection first. If the same load that caused the first protection trip is still connected the channel output will be disabled immediately after the protection programmed delay time expired. Therefore we also have to disconnect load or disable protection. The later method will be used in the command sequence that follows:

OUTP ON	First attempt to enable channel output
OUTP?
0	This attempt failed, the channel output remain disabled
OUTP:PROT:CLE	Channel protections reset
OUTP ON	Second to enable channel output
0	Channel output was enabled for a short time (100ms) and returns back to OFF state
CURR:PROT:TRIP?	Check OCP status
1	OCP has been tripped
OUTP:PROT:CLE	Reset channel protections once again
CURR:PROT:STAT OFF	Disable OCP
OUTP ON	Third attempt to enable channel output
OUTP?
1	Output is finally enabled
OUTP:MODE?
"CC"	Channel enters CC mode of operation

## Voltage and current calibration
For optimum calibration results the following condition are recommended:
- the calibration ambient temperature is stable and between 20 °C and 30 °C.
- ambient relative humidity is less than 80 %.
- Allow a one hour warm-up period before verification or calibration (use e.g. SYST:CHAN:INFO:ONT:LAST? or SYST:CPU:INFO:ONT:LAST? to get that info).
- Use short and thick cables to connect test setups.
Step	Commands	Description
1	*RST
2	SYST:RWL	Make sure that BB3 is in remote mode and cannot be unlock from local console (TFT display)
3	INST {CH1|CH2}; OUTP ON	Select the channel to be calibrated and enable the channel output.
4	VOLT:PROT:STAT OFF / CURR:PROT:STAT OFF / POW:PROT:STAT OFF	Disable if required the voltage, current and power protections.
5	CAL ON, "<password>"	BB3 enters calibration mode on the channel selected in step 1. Both voltage and current on the selected channel are set to the MINimum value. The VOLT? and CURR? commands can be optionally used here to test channel output values.
6		For voltage calibration, connect a voltmeter across the power module’s output terminals.
7	CAL:VOLT:LEV 1,0.15	Set the channel to the first calibration point and output voltage to 150 mV
8	CAL:VOLT 0.1455	Enter the reading you obtained from the external voltmeter.
9	CAL:VOLT:LEV 2,38	Set the channel to the second calibration point and output voltage 38 V
10	CAL:VOLT 39.292	Enter the reading you obtained from the voltmeter.
11		For current calibration, connect current monitoring resistor (shunt) across the output terminals and connect the ammeter across the shunt resistor. Its resistance has to be less then 5 Ω and rated for 25 W or more for measuring MAX current level.
12	CAL:CURR:RANG LOW	Set low current range (i.e. 0 – 50 mA).
13	CAL:CURR:LEV 1,0.0005	Set the channel to the first calibration point and output current to 0.5 mA.
14	CAL:CURR 0.000592	Enter the reading you obtained from the external ammeter.
15	CAL:CURR:LEV 2,0.048	Set the channel to the second calibration point and output current to 48 mA.
16	CAL:CURR 0.049863	Enter the reading you obtained from the ammeter.
17	CAL:CURR:RANG HIGH	Set high current range (i.e. 0 – 5 A).
18	CAL:CURR:LEV 1,0.005	Set the channel to the first calibration point and output current to 50 mA.
19	CAL:CURR 0.00603	Enter the reading you obtained from the external ammeter.
20	CAL:CURR:LEV 2,4.8	Set the channel to the second calibration point and output current to 4.8 A.
21	CAL:CURR 5.0731	Enter the reading you obtained from the ammeter.
22	CAL:REM "<string>"	Record calibration information such as next calibration due date for future reference. The calibration string may contain up to 32 characters. You don’t need to enter current date and time since that information will be recorded automatically.
23	CAL:SAVE	Save to non-volatile memory new calibration data.
24	CAL OFF, "<password>"	BB3 exit calibration mode. Both voltage and current on the selected channel are again set to the MINimum value.
25	SYST:REM	Enable local console unlock. Alternatively SYST:LOC can be executed to make local console enabled again.

## Working with profiles
The following command sequence could be used to store current set of parameters to the profile location 4 in the non-volatile memory:

MEM:STAT:VAL? 4	Check to see if profile selected location is empty
0
MEM:STAT:NAME? 4	We can also check that by querying profile location name
"--Not used--"
INST CH1 / VOLT?;:CURR?;:OUTP?	Examine currently programmed output values of the first channel
0.00;0.00;0
INST CH2 / VOLT?;:CURR?;:OUTP?	Examine currently programmed output values of the second channel
0.00;0.00;0
VOLT 12;:CURR 300mA / INST CH1 / VOLT 12;:CURR 300mA / OUTP 1;:OUTP 1, CH2	Reprogram both channel output values that will be stored as a new profile
*SAV 4	All profile parameters is now storing on the selected location
MEM:STAT:NAME? 4	Check profile name
""
MEM:STAT:NAME 4, "Dual 12V/300mA, Output ON"	Set the profile name (only ASCII characters are allowed!)
MEM:STAT:NAME? 4	Check the profile name once again
"Dual 12V/300ma, Output ON"

We can now turn the BB3 off (when it enters the standby mode) turn it on again and check some of the programmed parameters:

SYST:POW 0	The BB3 enters the standby mode
SYST:POW 1	Returns back from the standby mode. Please note that this command can be executed with the minimum of 5 seconds delay otherwise a -200,"Execution error" will be generated (you can check that with the SYST:ERR? command)
VOLT?;:CURR?;:OUTP?	Query programmed voltage, current and output state of the currently selected channel
0.00;0.00;0	Returned data indicate that previously saved values in profile number 4 were not used
*RCL 4 / VOLT?;:CURR?;:OUTP?	Recall parameters from desired location and execute query once again
12.00;0.30;1	The channel output values are now programmed using the selected profile

We can automate above mentioned process that channel profile parameters stored in non-volatile memory are using on power up. First we’ll check what is a current status of automatic recall and what profile will be used in the case of automatic recall:

MEM:STAT:REC:AUTO?	Query status of automatic profile recall during power on sequence
0	Automatic recall is turned off
MEM:STAT:REC:AUTO ON	Turn on automatic recall
MEM:STAT:REC:SEL?	Query which profile will be used when automatic recall is turned on
0	Selected profile was 0
MEM:STAT:REC:SEL 4	Change power on profile to 4
SYST:POW 0	Switch the BB3 to the standby mode once again
SYST:POW 1	Returns back from the standby mode. Again wait at least 5 seconds before enters this command
VOLT?;:CURR?;:OUTP?	Query programmed voltage, current and output state of the currently selected channel
12.00;0.30;1	The channel output values are programmed using the selected profile

## Get identification info and self-test results
The BB3’s identification information could be beneficial when more then one instrument are controlled. Additionally in the following example information about self-test will be queried:

*IDN?	Query identification string
Envox,EEZ BB3 (STM32),0000000,M1 0.2	BB3 with two different channels is identified, the first channel is 0-50V/3A and the second is 0-40V/5A. Serial number is 00001, and firmware version M1.0.93
*TST?	Execute self-test and query result
0	Self-test is passed
DIAG:TEST?	Query additional information about self-test
"2, EEPROM, installed, passed", / "2, SD card, installed, passed", / "2, Ethernet, installed, passed", / "2, RTC, installed, passed", / "2, DateTime, installed, passed", / "4, BP option, not installed, failed", / "2, Fan, installed, passed", / "2, AUX temp, installed, passed", / "2, CH1 temp, installed, passed", / "2, CH2 temp, installed, passed"

The self-test could be performed even when the BB3 is in the standby mode. We’ll first switch the BB3 into the standby mode. At the end of this example we are using additional diagnostic command that allows us to query information about channel’s ADC measurements.

SYST:POW 0	The BB3 enters the standby mode
DIAG:TEST?	Query additional information about self-test
"2, EEPROM, installed, passed", / "2, SD card, installed, passed", / "2, Ethernet, installed, passed", / "2, RTC, installed, passed", / "2, DateTime, installed, passed", / "4, BP option, not installed, failed", / "2, Fan, installed, passed", / "2, AUX temp, installed, passed", / "2, CH1 temp, installed, passed", / "2, CH2 temp, installed, passed"
SYST:POW 1	Returns back from the standby mode
DIAG:ADC?	Additional information about currently selected channel ADC inputs
"U_SET=12.02 V","U_MON=12.00 V","I_SET=0.30 A","I_MON=0.00 A"	U_SET and I_SET are measured values of the DAC outputs, U_MON and I_MON are actual output values. I_MON is 0 because no load is connected. A small difference between set and actual output voltage exists because calibration data are currently in use.

## Programming output voltage using the list of values
The BB3 comes with simple “arbitary waveform generator” functionality that can be accomplished using the LIST commands. The following example changes in the loop output voltage between five output values each half a second long while current is set to 3 A.

INST CH1	Select the channel that has to be programmed
VOLT:MODE LIST	Set voltage programming mode to the list of values
LIST:VOLT 5, 10, 20, 40, 0	Define sequence of output voltage
CURR:MODE LIST	Set current programming mode to the list of values.
LIST:CURR 3	Only one output current value is defined. Single value or number of values equivalent to other parameters (LIST:VOLT and LIST:DWEL) is allowed.
LIST:DWEL 0.5	Only one value for the duration of each step is chosen. Single value or number of values equivalent to other parameters (LIST:VOLT and LIST:CURR) is allowed.
LIST:COUN INF	Repeat continuously LIST sequence
TRIG:SOUR IMM	Define type of trigger. In this case the list execution will start immediately after INIT command is received.
INIT	Start the trigger.

Resulting output voltage and current waveform with connected load of 4.7 Ω is shown on the picture below.