## DISPlay
The DISPlay commands are used to set the display mode, turn on or off the front panel TFT display, select main page appearance, display and clear the text sent using a controller application. DISPlay is independent of, and does not modify, how data is returned to the controller application.

SCPI command	Description
DISPlay
:BRIGhtness {<value>}	Sets the intensity of the front panel TFT display
:DATA?	Reads screen image data
:VIEW {<mode>}	Sets the front panel TFT display main page appearance
[:WINdow]
[:STATe] {<bool>}	Sets the front panel TFT display state
:DIALog
[:OPEN] {<filename>}	Displays content of the resource file
:ACTIon? [<timeout>]	Returns information about selected item
:CLOSe	Closes last opened dialog window
:DATA {<name>}, {<type>}, [{<unit>}], {<value>}	Sets the value of the dialog data item
:DLOG [<filename>]	Opens DLOG viewer
:ERRor {<message>}	Displays message box and generates error beep tone
:INPUt? {<label>}, {<type>} [, <min>, <max>, <value>]	Displays entry form and waits for input on the front panel TFT display
:SELect? {<defaultSelection>, {<optionText>}[, ...]	Displays pop-up display with multiple radio buttons
:TEXT {<message>}	Displays a message on the front panel TFT display
:CLEar	Clear a message on the front panel TFT display

### DISPlay:BRIGhtness
Syntax	DISPlay:BRIGhtness {<value>} / DISPlay:BRIGhtness?
Description	Controls the intensity of the front panel TFT display. The range of the parameter is 1 to 20, where 20 is full intensity and 1 is fully blanked.
Parameters	Name	Type	Range	Default
	<value>	NR1	1 – 20	20
Return	This query returns set front panel's TFT display brightness value.
Usage example	DISP:BRIG? / 20
Related Commands	*RST

### DISPlay:DATA?
Syntax	DISPlay:DATA?
Description	This query reads screen image data. The image is formatted as a .png file. / Use the HCOPy[:IMMediate] or HCOPy:SDUMp[:IMMediate] command to capture screen image and save it as a file on the SD card.
Return	Screen image data is returned in the IEEE-488.2 # data block format (see Section 2.10).
Usage example	DISP:DATA? / #<length-digits><length><block>
Related Commands	HCOPy[:IMMediate] / HCOPy:SDUMp[:IMMediate] / MMEMory:FEED

### DISPlay:VIEW
Syntax	DISPlay:VIEW {<mode>} / DISPlay:VIEW?
Description	Use this command to set front panel TFT display main page appearance. The following modes are available: / 1 – Numerical / 2 – Bar graph horizontal / 3 – Bar graph vertical / 4 – YT view (scroll) / 5 – YT view (scan line)
Parameters	Name	Type	Range	Default
	<mode>	NR1|Discrete	1 – 5|DEFault	DEFault
Return	This query returns set front panel's TFT display main page appearance numeric value (NR1).
Usage example	DISP:VIEW 2
Related Commands	*RST

### DISPlay[:WINdow][:STATe]
Syntax	DISPlay[:WINdow][:STATe] {<bool>} / DISPlay[:WINdow][:STATe]?
Description	Turn the front panel TFT display off or on. When the display is turned off, outputs are not sent to the display and all indicators are disabled except the Event view indicator. The display state is automatically turned on when you return to the local mode. / Press and hold the display for about a second to return to the LOCal from the REMote control.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	ON|OFF|0|1	–
Return	DISPlay? query the front panel TFT display state. Returns 0 (OFF) or 1 (ON).
Usage example	DISP ON
Related Commands	SYSTem:LOCal / SYSTem:REMote

### DISPlay[:WINdow]:DIALog[:OPEN]
Syntax	DISPlay[:WINdow]:DIALog[:OPEN] {<filename>}
Description	Use this command to draw new page defined with resource file created in EEZ Studio and contains page widgets, actions and data items. The resource file has to be located on the SD card (see MMEMory subsystem) and the full path to the file must be specified. / The new displayed page can be closed (removed) by using the DISPlay[:WINdow]:DIALog:CLOSe to close the page.
Parameters	Name	Type	Range	Default
	<filename>	Quoted string	–	–
Usage example	DISP:DIAL "/Scripts/Diode Tester.res"
Related Commands	DISPlay[:WINdow]:DIALog:CLOSe / DISPlay[:WINdow]:DIALog:DATA

### DISPlay[:WINdow]:DIALog:ACTIon
Syntax	DISPlay[:WINdow]:DIALog:ACTIon? [<timeout>]
Description	This command waits for timeout (in seconds) to get information about what action item is selected on the currently displayed page as defined in resource file activated with the DISPlay[:WINdow]:DIALog[:OPEN] command. The default timeout value is 0 and command will wait until user select one of the action items.
Parameters	Name	Type	Range	Default
	<timeout>	NR1	–	0
Return	Returns name of the latest action item selected by user. 0 will be returned if currently opened page does not contain any action dialog. 1 is returned when timeout parameter is used and set time expired without user action.
Usage example	DISP:DIAL:ACTI? / "input_diode_name"
Related Commands	DISPlay[:WINdow]:DIALog[:OPEN]

### DISPlay[:WINdow]:DIALog:CLOSe
Syntax	DISPlay[:WINdow]:DIALog:CLOSe
Description	Use this command to close last dialog window opened with the DISPlay[:WINdow]:DIALog[:OPEN] command.
Usage example	DISP:DIAL:CLOS
Related Commands	DISPlay[:WINdow]:DIALog[:OPEN]

### DISPlay[:WINdow]:DIALog:DATA
Syntax	DISPlay[:WINdow]:DIALog:DATA {<name>}, {<type>}, [{<unit>}], {<value>}
Description	Use this command to set the value of the dialog data item defined as <name>. The list of possible data items is defined within the dialog resource file created in the EEZ Studio and specified with the DISPlay[:WINdow]:DIALog:OPEN command. / If <type> is FLOat then <unit> has to be specified, too.
Parameters	Name	Type	Range	Default
	<name>	Quoted string		–
	<type>	NR1|NR3|Quoted string	INTeger|FLOat|STRing	–
	<unit>	Discrete	UNKNown|VOLT| AMPEr|WATT|JOULe|SECOnd|OHM|FARAd|HERTz	–
Usage example	DISP:DIAL:DATA "test", INT
Related Commands	DISPlay[:WINdow]:DIALog:OPEN

### DISPlay[:WINdow]:DLOG
Syntax	DISPlay[:WINdow]:DLOG [<filename>]
Description	This command opens DLOG viewer and displays specified by filename. If filename is omitted it displays the last log file during or after recording is finished.
Usage example	DISP:DLOG "Recording/test_log.dlog"
Related Commands	INITiate:DLOG / INITiate:DLOG:TRACe

### DISPlay[:WINdow]:ERRor
Syntax	DISPlay[:WINdow]:ERRor {<message>}
Description	Use this command to display message box and generates error beep tone.
Usage example	DISP:ERR "Test message!"
Related Commands	DISPlay[:WINdow]:TEXT / SYSTem:BEEPer[:IMMediate]

### DISPlay[:WINdow]:INPut
Syntax	DISPlay[:WINdow]:INPut? {<label>}, {<type>} [, <min>, <max>, <value>]
Description	Use this command to define input field on the display that can returns a value of type <type> which can be TEXT | NUMBer | INTeger | MENU. The following variants are possible: / DISPlay[:WINdow]:INPut? {<label>}, TEXT, <min_length>, <max_length>, <value> / DISPlay[:WINdow]:INPut? {<label>}, NUMBer, <unit>, <min>, <max>, <value> / DISPlay[:WINdow]:INPut? {<label>}, INTeger, <min>, <max>, <value> / DISPlay[:WINdow]:INPut? {<label>}, MENU, <menu_type>, <menu_item>, [… <menu_item>] / <min> and <max> values should be defined for TEXT, NUMBer and INTeger. / For NUMBer and INTeger this is the minimum and maximum allowable value, and for STRing it is the minimum and maximum allowable text length. NUMBer should also be set to <unit> (see above which units are allowed). / If <type> is MENU then you should set <menu_type> (which can only be BUTTon) and up to max. 4 menu items (button texts) to be offered for selection. / A numeric keypad will be open for NUMBer and INTeger, and a full keypad will be open for TEXT. The default <label> (quoted string) will be displayed on the screen. / In all variants except MENU (where no need to be specified), the <value> parameter is the initial (or default) value.
Parameters	Name	Type	Range	Default
	<label>	Quoted string	–	–
	<type>	Discrete	TEXT|NUMBer|MENU
Return	This query returns value of defined type.
Usage example	DISP:INP? "Enter you text", TEXT, 2, 10, "Hello" / "Test"

### DISPlay[:WINdow]:SELect?
Syntax	DISPlay[:WINdow]:SELect? {<defaultSelection>, {<optionText>}[, ...]
Description	Display pop-up display with multiple radio buttons. The <defaultSelection> specify which of the presented option will be selected by default.
Parameters	Name	Type	Range	Default
	<defaultSelection>	NR1	1 to the total number of options	–
	<optionText>	Quoted string	–	–
Return	The query returns NR1 value of the selected option.
Usage example	Default option is 2 and query returns 3 when "Option 3" is selected: / DISP:SEL? 2, "Option 1", "Option 2", "Option 3" / 3
Related Commands	DISPlay[:WINdow]:DIALog:CLOSe / DISPlay[:WINdow]:DIALog:DATA / DISPlay[:WINdow]:INPut

### DISPlay[:WINdow]:TEXT
Syntax	DISPlay[:WINdow]:TEXT {<text message>} / DISPlay[:WINdow]:TEXT?
Description	Display a message on the front panel TFT display. The BB3 will display up to 32 characters in a message.
Parameters	Name	Type	Range	Default
	<text message>	Quoted string	–	–
Return	Query the message sent to the front panel TFT display and returns a quoted string.
Usage example	Send "Hello world" textual message: / DISP:TEXT "Hello world"
Related Commands	DISPlay[:WINdow]:TEXT:CLEar

### DISPlay[:WINdow]:TEXT:CLEar
Syntax	DISPlay[:WINdow]:TEXT:CLEar
Description	Clear the message displayed on the front panel TFT display.
Return	None
Usage example	DISP:TEXT:CLE
Related Commands	DISPlay[:WINdow]:TEXT