## INITiate
INITiate commands initialize the trigger system. This enables the trigger system to receive triggers.

SCPI command	Description
INITiate
[:IMMediate]	Completes one full trigger cycle
:DLOG {<filename>}	Initiates automated internal data log session
:TRACe {<filename>}
:CONTinuous {<bool>}	Enables/disables continuous transient triggers

### INITiate
Syntax	INITiate[:IMMediate]
Description	The INITiate command is used to initialize the trigger system of the BB3. This command completes one full trigger cycle when the trigger source is an IMMediate and initiates the trigger subsystem when the trigger source is BUS. / For example, when the IMMediate is selected as a trigger source, an INITiate command / immediately transfers the VOLTage:TRIGgered[:AMPLitude] and CURRent:TRIGgered[:AMPLitude] values to VOLTage[:LEVel][:IMMediate][:AMPLitude] and CURRent[:LEVel][:IMMediate][:AMPLitude] values. Any delay is ignored. / Execution of this command also affects bit 5 (Waiting for TRIGger) of the Operation Instrument Isummary register (see Section 3.3.2).
Usage example	Generate a trigger operation after 5 seconds: / TRIG:SOUR BUS / TRIG:DEL 5 / INIT / *TRG
Errors	-213,"Init ignored" / 307,"List lengths are not equivalent" / 308,"Cannot be changed while transient trigger is initiated" / 309,"Cannot initiate while in fixed mode"
Related Commands	*TRG / INSTrument:COUPle:TRIGger

### INITiate:DLOG
Syntax	INITiate:DLOG {<filename>}
Description	The command starts automated internal data logging session. All selected measurements defined with SENSe:DLOG:FUNCtion:CURRent, SENSe:DLOG:FUNCtion:POWer and SENSe:DLOG:FUNCtion:VOLTage commands will be recorded periodically in the specified log filename that is formatted as described below. The data logging session will last until time specified with the SENSe:DLOG:TIME not expired or is not interrupted prematurely using the e.g. ABORt:DLOG, SYSTem:REStart or *RST command. / Log file format / Data log structure allows recording of multiple values simultaneously that can be also viewed simultaneously on the same X-Y graph as in case of e.g. curve tracer when each Y-axis data set (column) represents one trace. This minimizes the required size of the data log file because the X-axis content is common for multiple values to be displayed on the Y-axis. / Data log file contains fixed header, flexible header and data section. / Fixed header / The fixed header includes the following information in little-endian format, i.e. the least significant byte (LSB) value is at the lowest address:
	Position	Type		Description
	0 – 7	ASCII		Always contains “EEZ-DLOG” text in ASCII format
	8 – 9	UINT16		File format version (e.g. 0x00 0x02)
	10 – 11	UINT16		Number of columns
	12 – 15	UINT32		File offset to data rows
	Flexible header / The flexible header contains meta information about data that follows. The number of fields defined in the flexible header is arbitrary (hence the name flexible header) formatted as follows: / <Field 1 length><Field 1 ID><Field 1 data> / <Field 2 length><Field 2 ID><Field 2 data> / <Field 3 length><Field 3 ID><Field 3 data> / ... / <Field N length><Field N ID><Field N data>
	Name	Length					Description
	Field length	UINT16					Total length of the field.
	Field ID	UINT8					Field unique identification number (see below the list of currently supported fields).
	Field data	For example if the <Field ID> is 14 or "X-axis label", then the field is of type String. Text string begins with UINT16 that contains its length, followed by length * UINT8 data.					Field data as specified below in the list of currently supported fields.

	List of currently supported fields:
	ID	Name			Data type			Description
	1	Comment			String			Data log user comment. Max. length is 128 characters.
	10	Unit			UINT8			0 – Unknown / 1 – Volt / 3 – Ampere / 6 – Watt / 8 – Second / 12 – Ohm / 16 – Hertz / 17 – Joule/Ws / 21 – Farad
	11	X-axis step			Float			X-axis resolution in units as specified in ID10. For example, if X-axis unit is 8 (seconds), when period of data logger is set to 20 ms this field will contains 0.02.
	12	X-axis range min value			Float			X-axis minimal or starting value which is 0 in most of the cases.
	13	X-axis range max value			Float			X-axis maximal or ending value.
	14	X-axis label			String			X-axis data description
	15	X-axis scale			UINT8			0 – Linear / 1 – Logarithmic
	30	Y-axis unit			UINT8 + UINT8			First UINT8 value specifies column number and second UINT8 contains Y-axis unit similar as X-axis unit specified under ID10.
	32	Y-axis range min value			UINT8 + Float			UINT8 value specifies column number followed by minimum recorded value.
	33	Y-axis range max value			UINT8 + Float			UINT8 value specifies column number followed by maximum recorded value. For example for uncoupled DCP405 module that will be 40 V and 5 A or for coupled in series or parallel 80 V or 10 A. This parameter is used by Auto-scale option in BB3 built-in log viewer.
	34	Y-axis label			UINT8 + String			UINT8 value specifies column number followed by Y-axis data description. This field is optional and if its value is not defined, the data log viewer (e.g. BB3 built-in viewer) could generate it from the channel number and unit symbol, for example: U1 for recorded voltage on the channel 1.
	35	Y-axis channel			UINT8 + UINT8			First UINT8 value specifies column number and second UINT8 contains instrument channel that is recorded
	36	Y-axis scale			UINT8 + UINT8			Similar to X-axis scale: first UINT8 value specifies column number and second UINT8 can contains: / 0 – Linear / 1 – Logarithmic
	50	Channel module type			UINT8 + UINT16			UINT8 value specifies channel number, and UINT16 contains module type ID as follows: / 220 – DCM220 / 405 – DCP405
	51	Channel revision			UINT8 + UINT16			UINT8 value specifies channel number, and UINT16 contains module version number, for example 0x0206 for “R2B6”.
	Logged data
	For the each X-axis step the series of Y values (columns) are recorded. All Y values are of type float. / <Y1 value><Y2 value>...<Yn value> / <Y1 value><Y2 value>...<Yn value> / ... / This command will automatically set the following fields: / X-axis unit = 8 / X-axis step = SENSe:DLOG:PERiod? / X-axis range min = 0 / X-axis range max = SENSe:DLOG:TIME? / For example when voltage and current are logged for channel 1 and 2 on two DCP405 modules, additional fields will be populated as follows: / Y1-axis unit = <1><1> / Y1-axis range min = 1 + SOUR1:VOLT? MIN / Y1-axis range max = 1 + SOUR1:VOLT? MAX / Y1-axis channel index = 1 / Y2-axis unit = <2><3> / Y2-axis range min = 2 + SOUR1:CURR? MIN / Y2-axis range max = 2 + SOUR1:CURR? MAX / Y2-axis channel index = 1 / Y3-axis unit = <3><1> / Y3-axis range min = 3 + SOUR2:VOLT? MIN / Y3-axis range max = 3 + SOUR2:VOLT? MAX / Y3-axis channel index = 2 / Y4-axis unit = <4><3> / Y4-axis range min = 4 + SOUR2:CURR? MIN / Y4-axis range max = 4 + SOUR2:CURR? MAX / Y4-axis channel index = 2 / Channel 1 module type = “405” / Channel 1 module revision = “207” / Channel 2 module type = “405” / Channel 2 module revision = “207”
Parameters	Name		Type			Range			Default
	<filename>		Quoted string			File name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters			–
Usage example	Start data logging of channel 1 voltage and current with resolution of 0.1 seconds that lasts 60 seconds: / SENS:DLOG:PER 0.1 / SENS:DLOG:TIME 60 / SENS:DLOG:FUNC:VOLT ON, CH1 / SENS:DLOG:FUNC:CURR ON, CH1 / INIT:DLOG "Recordings/test_log.dlog"
Related Commands	*RST / ABORt:DLOG / DISPlay[:WINdow]:DLOG / INITiate:DLOG:TRACe / MMEMory:INFOrmation? / SENSe:DLOG:FUNCtion:CURRent / SENSe:DLOG:FUNCtion:POWer / SENSe:DLOG:FUNCtion:VOLTage / SENSe:DLOG:PERiod / SENSe:DLOG:TIME

### INITiate:DLOG:TRACe
Syntax	INITiate:DLOG:TRACe {<filename>}
Description	The command starts internal data logging session. Single X-axis and multiple Y-axis are allowed as in case on the INITiate:DLOG:TRACe command, but offering more flexibility in defining recording parameters. The duration of data logging is not defined in advance and new data can be added until termination command is not executed e.g. ABORt:DLOG, SYSTem:REStart or *RST command.
Parameters	Name	Type	Range	Default
	<filename>	Quoted string	File name, either / (slash) or \ (backslash) can be used as the path separator. 1 to 255 characters	–
Usage example	Data logging is a four step process as follows: / 1. Define X- and Y-axis parameters. Number of Y-axis is arbitrary and two is used in this example. Data log file remark is not mandatory. / SENS:DLOG:TRAC:X:UNIT SECO / SENS:DLOG:TRAC:X:STEP 0.01 / SENS:DLOG:TRAC:X:RANG:MIN 0 / SENS:DLOG:TRAC:X:RANG:MAX 20 / SENS:DLOG:TRAC:X:LAB "t" / SENS:DLOG:TRAC:Y1:UNIT VOLT / SENS:DLOG:TRAC:Y1:LAB "U" / SENS:DLOG:TRAC:Y1:RANG:MIN 0 / SENS:DLOG:TRAC:Y1:RANG:MAX 40 / SENS:DLOG:TRAC:Y2:UNIT AMPE / SENS:DLOG:TRAC:Y2:LAB "I" / SENS:DLOG:TRAC:Y2:RANG:MIN 0 / SENS:DLOG:TRAC:Y2:RANG:MAX 5 / SENS:DLOG:TRAC:X:SCAL LIN / SENS:DLOG:TRAC:Y:SCAL LIN / 2. Define data log filename and comment/description. / SENS:DLOG:TRAC:REM "data log test" / INIT:DLOG:TRAC "Recordings/test_log.dlog" / 3. Add arbitrary number of recordings: / SENS:DLOG:TRAC:DATA 10.1,2.55 / SENS:DLOG:TRAC:DATA 12.0,2.66 / SENS:DLOG:TRAC:DATA 16.34,3.63 / … / 4. Close data logging session. / ABOR:DLOG
Related Commands	*RST / ABORt:DLOG / DISPlay[:WINdow]:DLOG / INITiate:DLOG / MMEMory:INFOrmation? / SENSe:DLOG:TRACe:REMark / SYSTem:REStart

### INITiate:CONTinuous
Syntax	INITiate:CONTinuous {<bool>} / INITiate:CONTinuous?
Description	This command is used to select whether the trigger system is continuously initiated or not. With CONTinuous set to OFF, the trigger system remain in the IDLE state until CONTinuous is set to ON or INITiate:IMMediate is received. / Once CONTinuous is set to ON, the trigger system will be initiated and exit the IDLE state. On completion of each trigger cycle, with CONTinuous ON, the trigger system immediately commence another trigger cycle without entering the IDLE state. / When INITiate:CONTinuous is set to OFF, the current trigger cycle will be completed before entering the IDLE state. The return to IDLE also occur as the result of an ABORt or *RST command. / The ABORt command force the trigger system to the IDLE state; however, the value of INITiate:CONTinuous is unaffected. / If INITiate:CONTinuous was set to ON prior to receiving ABORt, it remains ON and the trigger system immediately exit the IDLE state.
Parameters	Name	Type	Range	Default
	<bool>	Boolean	OFF|ON|0|1	OFF
Return	The query command returns 0 if continuous transients are disabled (OFF), and 1 if continuous transients are enabled (ON).
Usage example	INIT:CONT ON
Related Commands	*RST / ABORt